ATTACHING THE TERMINAL

The MAX180 evaluation board is to be attached to a dumb  terminal

or a personal computer running a terminal emulation program.  If

the serial port on the terminal or personal computer has a 25-pin

connector a standard straight through cable can be used to attach

the  board.  An adapter is needed for use with 9-pin  connectors.

See the manual for more information.

OPERATING THE SOFTWARE

Maxim supplies the SERCOMM terminal emulation program for working

with this evaluation kit.   The SERCOMM.EXE program on this disk-

ette  may  be  copied to a hard disk drive or  started  from  the

floppy disk.  The program displays the active communication  port

number  and  baud rate on the bottom line of  the  display.   The

active  port  or baud rate may be changed by entering  ctrl-T  or

ctrl-R respectively.  SERCOMM will attempt to reset the board and

send a space character whenever the port or baud rate is changed.

The  evaluation  board will also work  with  commercial  terminal

emulator  programs such as XTALK or PROCOMM.  Before  turning  on

the  evaluation  board  you should start  the  terminal  emulator

program  and set it for 8-bit, no parity, full duplex  operation.

The  baud rate may be set  to any standard baud rate from 300  to

9600 baud.

When power is applied to the board, the firmware does some   ini-

tialization  and  then enters an automatic  baud  rate  detection

loop.  The autobaud routine examines the first character received

to  determine the incoming baud rate.  In order for this  routine

to operate correctly,

****************************************************************

****THE FIRST CHARACTER SENT AFTER POWER UP MUST BE A SPACE.****

****************************************************************

If the first character received by the board was not a space  the

autobaud routine will miscalculate the baud rate and all communi-

cations will be garbled.  The only way to correct this  condition

is a hardware reset (such as a power off/on cycle).

If  the  terminal occasionally misses a character from  the  data

stream, try setting the system to a slower baud rate.

After communications has been established a copyright banner is

displayed followed by a descriptive command prompt.

THE COMMAND PROMPT

The command prompt takes the form

MAX180:current OUTPUT format:current INPUT mode:CHANNEL number >

where CURRENT OUTPUT FORMAT is

HEX, DEC or BIN

and CURRENT INPUT MODE is either

UNI or BIP

and the CURRENT CHANNEL NUMBER is in the following set:

{0,1,2,3,4,5,6,7}       For single ended operation

or

{0/1,2/3,4/5,6/7}       For differential operation

{1/0,3/2,5/4,7/6}

In differential mode, the first channel given is the analog  high

input.   That  is,  the output will read positive  if  the  input

voltage  to  this channel is higher than the other  (analog  low)

channel.  All possible differential mode channel pairs are listed

above.

You  can determine the current operating mode of the  monitor  by

examining the prompt.  For instance, if the prompt reads:

MAX180:HEX:BIP:0/1 >

The  MAX180 is encoding data using signed 2's  complement 11-bit

coding and the output format would be in ASCII hexadecimal.   The

input is in differential mode using the channel 0-1 analog  input

pair.  The output will give a positive reading if channel 0 is

greater than channel 1.

The  command input buffer will store a maximum of  8  characters.

An error message ("?")  will be returned  if more then 8  charac-

ters are typed without an ENTER ( a Carriage Return).   You  will

also receive the error message for an unknown command.

THE CONVERTER COMMANDS

HELP  is available with the help command, "?".  The  help  screen

lists  all  of the commands displayed in  functional  groups,  as

follows:

----------------------------------------------------------------

? - This HELP screen.

Display Output Mode Selection Commands:

H - Hexadecimal   Format

D - ASCII Decimal Format

^B - Binary Values (12 bits/2bytes, MSByte 1st)

MAX180 Conversion Mode Selection Commands:

B - Bipolar Mode

U - Unipolar Mode

MAKE CONVERSION Commands:

M - Make a single conversion and display result.

A NULL command (i.e. just CR) also makes a conversion.

C - Continuously convert and display until key is pressed.

Channel Select Commands:

0...7 - Select Channel Number

DI - Differential Input Mode

SE - Single-Ended Input Mode

RAM Related Commands:

R - RAM test

S - Store conversions in RAM buffer until full.

O - Output the values stored in RAM buffer.

----------------------------------------------------------------

The  Display Output Mode Selection Commands are used to  set  the

display  format  of MAX180 conversion results.  When  the  MAX180

makes  a  conversion the result is always 12 bits long.   The  12

bits of data are presented as follows:

H - Hexadecimal Format

The 12 bit value is converted into ASCII HEXADECIMAL format.  For

unipolar  mode,  the "000" to "FFF" codes represent  0  to  5.000

volts.   For bipolar mode, the input range is -2.5 volts to  +2.5

volts. For -2.50 volts the output is "800".

D - Decimal Format

The  12  bit value is converted into ASCII DECIMAL  format.   The

dynamic range is dependent on the current input mode, as follows:

INPUT MODE             DYNAMIC RANGE

Unipolar (UNI)           0 ... 4.999

Bipolar  (BIP)      -2.500 ... 2.499

^B - Binary Values (12 bits/2bytes, MSByte 1st)

DO  NOT USE THIS COMMAND TO SEND DATA TO THE DISPLAY SCREEN.  The

binary  values  will have unpredictable control effects  on  your

system.   This mode is used to capture data to a disk  file  when

using commercial terminal emulators.

The  binary values are returned as raw binary data in   two  byte

packets.   The first byte has four leading zeros followed by  the

most  significant four bits of the conversion.  The  second  byte

holds the least significant 8 bits of the conversion.

The  Conversion  Mode  Selection Commands are used  to  tell  the

MAX180 which Data Transfer Function it should use to present  the

result of the conversion.

B - Bipolar Mode

The MAX180 yields a signed 11-bit signed two's complement number.

The input range is -2.500 volts to +2.499 volts.

U - Unipolar Mode

This  is  the default operating mode of the  system.   The  input

voltage range is from 0 to +5.00 Volts.

The  following  commands are used to select  the  current  analog

input configuration:

{0..7} - Select the current channel.

The  entered channel number becomes the "current  channel".   The

prompt  will change to the given channel number.  If the  current

input  mode  is  DIFFERENTIAL the entered channel  will  be   the

analog high input of the channel pair number.  That is, if  chan-

nel 3 is selected, the prompt will indicate channel "3/2" as  the

current channel.

DI - Select Differential Input Mode

Analog  input  channels are paired.  Instead  of  8  Single-Ended

analog channels there are four Differential analog channels.

SE - Single Ended Input Mode

Analog input channels are Single Ended, that is the input voltage

is measured with respect to analog ground rather then another  of

the MAX180's input channels.

R - RAM test

The  firmware will  continuously fill the RAM with the  following

bit patterns and then reads the data back to verify RAM  integri-

ty.

BINARY         HEX

-------------------

00000000       00

11111111       FF

01010101       55

10101010       AA

The RAM test runs until the any key is pressed.  The results  are

displayed as errors are detected or at the beginning of each test

pattern run.

S - Store conversions in RAM buffer until full.

4096 conversion are made in the current mode (UNI or BIP) and the

results are stored in the RAM as raw 12-bit data.

O - Output the values stored in RAM buffer.

The  current  contents of the RAM are formatted per  the  current

output  mode  (HEX, DEC or BIN) and the results are sent  to  the

terminal.

Keep  the following in mind when outputting the contents  of  the

RAM  buffer.  If you are using DECIMAL formatting do  NOT  change

between  UNI  and  BIP input modes between the time  you  take  a

sample  (S) and do the output (O) because the output coding  will

not match the input coding.

#########################################################################################