| # Agilent CSA N1996A Spectrum Analyzer Upgrades |  |  |
| --- | --- | --- |

| Most of the Agilent CSA options are available for retrofit. The upgrades are normally either a Enabled Hardware Option or Measurement Personality. Some of the measurement personalities are application specific and may require additional hardware or accessories as well. Retrofit kits requiring the installation of hardware will contain the assemblies to be installed as well as the installation instructions. In many cases, a license key will be required to activate the measurement personality or hardware option. Below you will find information on the ordering requirements for: - Internal Hardware Upgrades - External Hardware Upgrades - Measurement Personality & Firmware update - Additional Documentation Options ## Hardware Upgrades: - Many hardware upgrades do require some level of adjustments and performance verification testing once the installation has been completed, many of which will require additional test equipment. For a complete list of all required adjustments and performance verification testing see the installation note for each option upgrade desired. - All adjustments will require the use of the Agilent CSA Performance Verification & Adjustment Software, which will also automate the required performance verification testing. - The cost of installation must also be considered if the option is to be installed at an Agilent Technologies Service Center. - Many upgrades will require the redemption and installation of a license key to be activated. Once a redemption certificate is received they can be redeemed at the Agilent Software License Redemption Web Site Agilent CSA Internal Hardware Upgrades. Option; Description; Upgrade Available?; Order As; Installation Note; Upgrade Contents1; Minimum Firmware; Adjustments Required2; Additional Requirements; Incompatible With; Installed By; Installation Time. 506; Frequency Range 100KHz to 6GHz; No3; n/a; n/a; n/a; n/a; n/a; n/a; n/a; n/a; n/a. BAT; Battery Pack (2 each); Yes; N1996AK-BAT (inc. 2) or 1420-0891 (inc. 1); n/a; H; Any; No; None; None; Customer; n/a. PO3; Preamplifier, 3GHz; Yes; N1996AK-P03; Available; L; Any; Yes6; 503; 506; Customer; 2.0 Hours. PO6; Preamplifier, 6GHz; Yes; N1996AK-P06; Available; L; Any; Yes6; 506; 503; Customer; 2.0 Hours. TG3; Tracking Generator, 3GHz; No4; N/A; Available; L; Any; Yes6; 503 & N8995S; 506; Customer; 2.0 Hours. TG6; Tracking Generator, 6GHz; No4; N/A; Available; L; Any; Yes6; 506 & N8995S; 503; Customer; 2.0 Hours. AFM; AM/FM Tune & Listen; Yes; N1996AK-AFM; Available; L; >=A.02.00; No; None; None; Customer; 0.25 Hours. Agilent CSA External Hardware Upgrade. Option; Description; Upgrade Available?; Order As; Installation Note; Upgradde Contents1; Minimum Firmware; Adjustments Required2; Additional Requirements; Compatible With; Installed By; Installation Time. 1CM; Rackmount Kit with Handles; Yes; N1996AK-1CM or N1996-60028; Available; H; Any; No; None; 1CP & SCC; Customer; 0.5 Hours. 1CP; Rackmount Kit w/o Handles; Yes; N1996AK-1CP or N1996-60029; Available; H; Any; No; None; 1CM & SCC; Customer; 0.5 Hours. BCG; External Battery Charger; Yes; N1996AK-BCG or 1950-4276; n/a; H; Any; No; BAT; None; Customer; n/a. HTC; Hard Transit Case; Yes; N1996AK-HTC or N1996-80013; n/a; H; Any; No; None; 1CM & 1CP; Customer; n/a. SCC; Soft Carry Case; Yes; N1996AK-SCC or N1996-80010; n/a; H; Any; No; None; 1CM & 1CP; Customer; n/a. SRK; Stimulus Response Calibration Kit; Yes; N1996AK-SRK or N1996-60022; n/a; H; Any; No; TG3 or TG6; None; Customer; n/a. --; Benchtop Bottom Feet Kit; Yes; 5000-0913; Available; H; Any; No; None; None; Customer; 0.25 Hours. --; Benchtop Rear Feet Kit; Yes; N1996-60021; H; Any; No; None; None; Customer; 0.25 Hours. ## Measurement Personalities: - Measurement Personality Options may have specific instrument firmware requirements. - All Measurement Personality upgrades require the redemption and installation of a license key to be activated. Once a redemption certificate is received they can be redeemed at the Agilent Software License Redemption Web Site - While all Measurement Personality upgrades only require that a license key and software be installed before the personality will function, there may also be additional hardware requirements. Agiletn CSA Measurement Personalities. Option; Description; Upgrade Available?; Order As; Upgrade Contents1; Minimum Firmware; Adjustment Required2; Additional Requirements; Incompatible With; Installed by; Installation Time. 271; Spectrogram Measurements; Yes; N1996AK-271; L; Any; No; None; None; Customer; 0.5 Hours. N8995A; Stimulus Response Measurement Suite; No5; N/A; L; Any; No; TG3 or TG6; None; Customer; 0.5 Hours. N8995A-SR3; Stimulus Response Measurement Suite, 3 GHz; Yes; N8995A-SR3; L; >=A.02.00; Yes6; None; None; Customer; 0.25 Hours. N8995A-SR6; Stimulus Response Measurement Suite, 6 GHz; Yes; N8995A-SR6; L; >=A.02.00; Yes6; None; None; Customer; 0.25 Hours. N8996A; AM/FM Modulation Analysis Suite; Yes; N8996A-1FP; L; >=A.02.00; No; None; None; Customer; 0.25 Hours. Agilent CSA Additional Documentation Options. Option; Description; Upgrade Available?; Order As; Upgrade Contents1; Minimum Firmware; Adjustments Required2; Aditional Requirements; Incompatible With; Installed by; Installation Time. OBW; Service Documentation; Yes; N1996AK-0BW; n/a; n/a; n/a; n/a; n/a; n/a; n/a. - H = Hardware, S = Software, L = License Key - All Adjustments require Performance Varification & Adjustment Software and specified test equipment. - Time of purchase only option. - Option TG3 has been replaced by N8995A-SR3. Option TG6 has been replaced by N8995A-SR6. - N8995A has been replaced by N8995A-SR3 and N8995A-SR6. - At the time of manufacture the hardware related to this option was fully adjusted and the option performance was verified to be within its warranted specifications. Within one year of the initial calibration date of the instrument this option is fully calibrated with no further adjustment or verification testing. |  | Weitere Informationen Technischer Support. Bibliothek. |
| --- | --- | --- |