## Address Map
Address Map W2000A Series
WELEC factory settings - depending on Altera reference design
hex address |  |  | Description
| offset | sector
0x | 0000 | 0000 | NIOS CPU
AM29LV065DU flash memory 128 x 64KB = 8 MB
0x | 0000 | 0000 | Flash base address
0x | 0000 | 0001 | first accessable address
0x | 0004 | 0000 | Bootloader
0x | 0004 | 0100 | Program start
0x | 000D | FFFF | Program end
0x | 000E | 0000 | not used
0x | 000F | 0000 | Confiig and storing area
0x | 0080 | 0000 | Flash end
IS61LV51216-8TL(512K x 16) SRAM 2MB
0x | 0080 | 0000 | RAM base address
0x | 0080 | 0001
0x | 00A0 | 0000 | RAM end
Peripherals
0x | 00A0 | 0000 | ADC 1 base address
0x | 0130 | 0200 | power on boot base address

## Flash
Flash Memory Map W2000A Series
version 1.2.BF.1.6
hex address |  |  | Description
| offset | sector
0x | 0000 | 0000 | NIOS CPU
AM29LV065DU flash memory 128 x 64KB = 8 MB
0x | 0000 | 0000 | Flash base address
0x | 0000 | 0001 | first accessable address
not used
|  |  | not used | reserved by OSOZ
0x | 0002 | 3CE1 | not used | reserved by OSOZ
|  |  | not used | reserved by OSOZ
0x | 0002 | 3D81 | not used | reserved by OSOZ
|  |  | not used | reserved by OSOZ
0x | 0003 | FFFF | not used | reserved by OSOZ
0x | 0004 | 0000 | Bootloader -> copies flash to RAM
0x | 0004 | 0100 | Program start
0x
0x
0x
0x
0x | 000D | FFFF | Program end
0x | 000E | 0000 | not used
0x | 000F | 0000 | protected_Config_Flash | 64 KB beginn config and storing area
0x | 0010 | 0000 | SAVE_SIGNAL_NR_A01 | 64 KB
0x | 0011 | 0000 | SAVE_SIGNAL_NR_A02 | 64 KB
0x | 0012 | 0000 | SAVE_SIGNAL_NR_A03 | 64 KB
0x | 0013 | 0000 | SAVE_SIGNAL_NR_A04 | 64 KB
0x | 0014 | 0000 | SAVE_SIGNAL_NR_A05 | 64 KB
0x | 0015 | 0000 | SAVE_SIGNAL_NR_A06 | 64 KB
0x | 0016 | 0000 | SAVE_SIGNAL_NR_A07 | 64 KB
0x | 0017 | 0000 | SAVE_SIGNAL_NR_A08 | 64 KB
0x | 0018 | 0000 | SAVE_SIGNAL_NR_A09 | 64 KB
0x | 0019 | 0000 | SAVE_SIGNAL_NR_A10 | 64 KB
0x | 001A | 0000 | SAVE_SIGNAL_NR_A11 | 64 KB
0x | 001B | 0000 | SAVE_SIGNAL_NR_A12 | 64 KB
0x | 001C | 0000 | SAVE_SIGNAL_NR_A13 | 64 KB
0x | 001D | 0000 | SAVE_SIGNAL_NR_A14 | 64 KB
0x | 001E | 0000 | SAVE_SIGNAL_NR_A15 | 64 KB
0x | 001F | 0000 | SAVE_SIGNAL_NR_A16 | 64 KB
0x | 0020 | 0000 | SAVE_SIGNAL_NR_B01 | 64 KB
0x | 0021 | 0000 | SAVE_SIGNAL_NR_B02 | 64 KB
0x | 0022 | 0000 | SAVE_SIGNAL_NR_B03 | 64 KB
0x | 0023 | 0000 | SAVE_SIGNAL_NR_B04 | 64 KB
0x | 0024 | 0000 | SAVE_SIGNAL_NR_B05 | 64 KB
0x | 0025 | 0000 | SAVE_SIGNAL_NR_B06 | 64 KB
0x | 0026 | 0000 | SAVE_SIGNAL_NR_B07 | 64 KB
0x | 0027 | 0000 | SAVE_SIGNAL_NR_B08 | 64 KB
0x | 0028 | 0000 | SAVE_SIGNAL_NR_B09 | 64 KB
0x | 0029 | 0000 | SAVE_SIGNAL_NR_B10 | 64 KB
0x | 002A | 0000 | SAVE_SIGNAL_NR_B11 | 64 KB
0x | 002B | 0000 | SAVE_SIGNAL_NR_B12 | 64 KB
0x | 002C | 0000 | SAVE_SIGNAL_NR_B13 | 64 KB
0x | 002D | 0000 | SAVE_SIGNAL_NR_B14 | 64 KB
0x | 002E | 0000 | SAVE_SIGNAL_NR_B15 | 64 KB
0x | 002F | 0000 | SAVE_SIGNAL_NR_B16 | 64 KB
0x | 0030 | 0000 | SAVE_SIGNAL_NR_C01 | 64 KB
0x | 0031 | 0000 | SAVE_SIGNAL_NR_C02 | 64 KB
0x | 0032 | 0000 | SAVE_SIGNAL_NR_C03 | 64 KB
0x | 0033 | 0000 | SAVE_SIGNAL_NR_C04 | 64 KB
0x | 0034 | 0000 | SAVE_SIGNAL_NR_C05 | 64 KB
0x | 0035 | 0000 | SAVE_SIGNAL_NR_C06 | 64 KB
0x | 0036 | 0000 | SAVE_SIGNAL_NR_C07 | 64 KB
0x | 0037 | 0000 | SAVE_SIGNAL_NR_C08 | 64 KB
0x | 0038 | 0000 | SAVE_SIGNAL_NR_C09 | 64 KB
0x | 0039 | 0000 | SAVE_SIGNAL_NR_C10 | 64 KB
0x | 003A | 0000 | SAVE_SIGNAL_NR_C11 | 64 KB
0x | 003B | 0000 | SAVE_SIGNAL_NR_C12 | 64 KB
0x | 003C | 0000 | SAVE_SIGNAL_NR_C13 | 64 KB
0x | 003D | 0000 | SAVE_SIGNAL_NR_C14 | 64 KB
0x | 003E | 0000 | SAVE_SIGNAL_NR_C15 | 64 KB
0x | 003F | 0000 | SAVE_SIGNAL_NR_C16 | 64 KB
0x | 0040 | 0000 | SAVE_SIGNAL_NR_D01 | 64 KB
0x | 0041 | 0000 | SAVE_SIGNAL_NR_D02 | 64 KB
0x | 0042 | 0000 | SAVE_SIGNAL_NR_D03 | 64 KB
0x | 0043 | 0000 | SAVE_SIGNAL_NR_D04 | 64 KB
0x | 0044 | 0000 | SAVE_SIGNAL_NR_D05 | 64 KB
0x | 0045 | 0000 | SAVE_SIGNAL_NR_D06 | 64 KB
0x | 0046 | 0000 | SAVE_SIGNAL_NR_D07 | 64 KB
0x | 0047 | 0000 | SAVE_SIGNAL_NR_D08 | 64 KB
0x | 0048 | 0000 | SAVE_SIGNAL_NR_D09 | 64 KB
0x | 0049 | 0000 | SAVE_SIGNAL_NR_D10 | 64 KB
0x | 004A | 0000 | SAVE_SIGNAL_NR_D11 | 64 KB
0x | 004B | 0000 | SAVE_SIGNAL_NR_D12 | 64 KB
0x | 004C | 0000 | SAVE_SIGNAL_NR_D13 | 64 KB
0x | 004D | 0000 | SAVE_SIGNAL_NR_D14 | 64 KB
0x | 004E | 0000 | SAVE_SIGNAL_NR_D15 | 64 KB
0x | 004F | 0000 | SAVE_SIGNAL_NR_D16 | 64 KB
0x | 0050 | 0000 | SAVE_SIGNAL_NR_E01 | 64 KB
0x | 0051 | 0000 | SAVE_SIGNAL_NR_E02 | 64 KB
0x | 0052 | 0000 | SAVE_SIGNAL_NR_E03 | 64 KB
0x | 0053 | 0000 | SAVE_SIGNAL_NR_E04 | 64 KB
0x | 0054 | 0000 | SAVE_SIGNAL_NR_E05 | 64 KB
0x | 0055 | 0000 | SAVE_SIGNAL_NR_E06 | 64 KB
0x | 0056 | 0000 | SAVE_SIGNAL_NR_E07 | 64 KB
0x | 0057 | 0000 | SAVE_SIGNAL_NR_E08 | 64 KB
0x | 0058 | 0000 | SAVE_SIGNAL_NR_E09 | 64 KB
0x | 0059 | 0000 | SAVE_SIGNAL_NR_E10 | 64 KB
0x | 005A | 0000 | SAVE_SIGNAL_NR_E11 | 64 KB
0x | 005B | 0000 | SAVE_SIGNAL_NR_E12 | 64 KB
0x | 005C | 0000 | SAVE_SIGNAL_NR_E13 | 64 KB
0x | 005D | 0000 | SAVE_SIGNAL_NR_E14 | 64 KB
0x | 005E | 0000 | SAVE_SIGNAL_NR_E15 | 64 KB
0x | 005F | 0000 | SAVE_SIGNAL_NR_E16 | 64 KB
0x | 0060 | 0000 | unused | 64 KB reserved by OSOZ
0x | 0061 | 0000 | unused | 64 KB reserved by OSOZ
0x | 0062 | 0000 | unused | 64 KB reserved by OSOZ
0x | 0063 | 0000 | unused | 64 KB reserved by OSOZ
0x | 0064 | 0000 | unused | 64 KB reserved by OSOZ
0x | 0065 | 0000 | unused | 64 KB reserved by OSOZ
0x | 0066 | 0000 | unused | 64 KB reserved by OSOZ
0x | 0067 | 0000 | unused | 64 KB reserved by OSOZ
0x | 0068 | 0000 | unused | 64 KB reserved by OSOZ
0x | 0069 | 0000 | unused | 64 KB reserved by OSOZ
0x | 006A | 0000 | unused | 64 KB reserved by OSOZ
0x | 006B | 0000 | unused | 64 KB reserved by OSOZ
0x | 006C | 0000 | WTAG_Logo_Half_Flash | 64 KB
0x | 006D | 0000 | WTAG_Logo | 64 KB
0x | 006E | 0000 | unused | 64 KB
0x | 006F | 0000 | unused | 64 KB
0x | 0070 | 0000 | unused | 64 KB
0x | 0071 | 0000 | unused | 64 KB
0x | 0072 | 0000 | unused | 64 KB
0x | 0073 | 0000 | unused | 64 KB
0x | 0074 | 0000 | unused | 64 KB
0x | 0075 | 0000 | unused | 64 KB
0x | 0076 | 0000 | unused | 64 KB
0x | 0077 | 0000 | unused | 64 KB
0x | 0078 | 0000 | TRIGOTAB01 | 64 KB sin/cos tables, first window table
0x | 0079 | 0000 | TRIGOTAB02 | 64 KB 4 window tables with 4096 x 4byte (int)
0x | 007A | 0000 | TRIGOTAB03 | 64 KB 4 window tables with 4096 x 4byte (int)
0x | 007B | 0000 | unused | 64 KB
0x | 007C | 0000 | CONFIG_FLASH | 64 KB actual configuration
0x | 007D | 0000 | CONFIG_BACKUP | 64 KB copy of configuration sector
0x | 007E | 0000 | CALIBRATION | 64 KB calibration data, hardware setting
0x | 007F | 0000 | unused | 64 KB -> old config sector
0x | 007F | FFFF | Flash end

## SRAM
SRAM Memory Map W2000A Series
version 1.2.BF.1.6
hex address |  |  | Description
| offset | sector
0x | 0000 | 0000 | NIOS CPU
IS61LV51216-8TL(512K x 16) SRAM 2MB
0x | 0080 | 0000 | RAM base address
0x | 0080 | 0001
0x | 008C | A5F4 | Buffer_Mark1Plane
0x | 008D | 3BF4 | Buffer_UI2Plane
0x | 008D | D1F4 | not used
0x | 008D | D320 | SIGNAL_Histo | QM calculation
0x | 008D | DD20 | SIGNAL1_Intpolate | shared memory for
0x | 008D | E1D0 | SIGNAL2_Intpolate | interpolation and USTB buffer
0x | 008D | E680 | SIGNAL3_Intpolate
0x | 008D | EB30 | SIGNAL4_Intpolate
0x | 008D | EFE0 | SIGNALM_Intpolate
0x | 008D | F490 | SIGNAL1_Intpolate_delayed | shared memory for
0x | 008D | F940 | SIGNAL2_Intpolate_delayed | interpolation and USTB buffer
0x | 008D | FDF0 | SIGNAL3_Intpolate_delayed
0x | 008E | 02A0 | SIGNAL4_Intpolate_delayed
0x | 008E | 0750 | SIGNALM_Intpolate_delayed
0x | 008E | 0C00 | SIGNAL1_Buffer | size 16Kb (0x4000) → overlay/FFT
0x | 008E | 4C00 | SIGNAL2_Buffer | size 16Kb (0x4000) → overlay
0x | 008E | 8C00 | SIGNAL3_Buffer | size 16Kb (0x4000) → overlay/FFT
0x | 008E | CC00 | SIGNAL4_Buffer | size 16Kb (0x4000) → overlay
0x | 008F | 0C00 | not used
0x | 008F | 4C00 | readout_sigbuf → not used anymore | size 16Kb (0x4000)
0x | 008F | 8C00 | FIR_Buffer / SHIFT_OFFS1 / ROLL_OFFS1 | USTB buffer
0x | 008F | CC00 | SHIFT_OFFS2
0x | 0090 | 0C00 | ROLL_OFFS2 | USTB 32Kb buffer
0x | 0090 | BC00 | ROLL_OFFS3 / Filter_1_Buffer
0x | 0090 | CC00 | Filter_2_Buffer | size 16Kb (0x4000)
0x | 0091 | 0C00 | ROLL_OFFS4 / Filter_3_Buffer | size 16Kb (0x4000)
0x | 0091 | 4C00 | Filter_4_Buffer | size 16Kb (0x4000)
0x | 0091 | BC00 | CopyBuffer | size 16Kb (0x4000)
0x | 0091 | CC00 | SIGNAL1 | size 16Kb (0x4000)
0x | 0092 | 0C00 | SIGNAL2 | size 16Kb (0x4000)
0x | 0092 | 4C00 | SIGNAL3 | size 16Kb (0x4000)
0x | 0092 | 8C00 | SIGNAL4 | size 16Kb (0x4000)
0x | 0092 | CC00 | SIGNALM | size 16Kb (0x4000)
0x | 0093 | 0C00 | Buffer_PlaneM / v_Buffer_PlaneM (long) | 38400b = 80b*480 = 8b*640*480
0x | 0093 | A200 | Buffer_Plane1 / v_Buffer_Plane1 (long)
0x | 0094 | 3800 | Buffer_Plane2 / v_Buffer_Plane2 (long)
0x | 0094 | CE00 | Buffer_Plane3 / v_Buffer_Plane3 (long)
0x | 0095 | 6400 | Buffer_Plane4 / v_Buffer_Plane3 (long)
0x | 0095 | FA00 | Planes | display planes base address
0x | 0095 | FA00 | UI_Plane1
0x | 0096 | 9000 | UI_Plane2
0x | 0097 | 2600 | UI_Plane3
0x | 0097 | BC00 | UI_Plane4
0x | 0098 | 5200 | UI_Plane5
0x | 0098 | E800 | Channel_Plane1 / v_Channel_Plane1 (long)
0x | 0099 | 7E00 | Channel_Plane2 / v_Channel_Plane2 (long)
0x | 009A | 1400 | Channel_Plane3 / v_Channel_Plane3 (long)
0x | 009A | AA00 | Channel_Plane4 / v_Channel_Plane4 (long)
0x | 009B | 4000 | Channel_Math_Plane / v_Channel_Math_Plane (long)
0x | 009B | D600 | Copy_Plane1
0x | 009C | 6C00 | Copy_Plane2
0x | 009D | 0200 | Copy_Plane3
0x | 009D | 9800 | Marker_Plane1
0x | 009E | 2E00 | Marker_Plane2
0x | 009E | C400 | Grid_Plane
0x | 009F | 5A00 | not used
0x | 009F | F000 | not used
0x
0x | 009F | FEFF | end program/data memory
0x | 009F | FF00 | stack top | vector table (size 0x00000100)
0x |  |  | stack
0x | 009F | FFFF | RAM end

## Peripherals
Peripheral Map W2000A Series
WELEC factory settings - depending on Altera reference design
hex address |  |  | Description
| offset | sector
0x | 0000 | 0000 | NIOS CPU
Peripherals
0x | 00A0 | 0000 | ADC 1
0x | 00A0 | 0010 | ADC 2
0x | 00A0 | 0020 | ADC 3
0x | 00A0 | 0030 | ADC 4
0x | 00A0 | 0040 | Watchdog | ALTERA AVALON PIO
0x | 0100 | 0000 | boot ROM base address | ALTERA AVALON ONCHIP MEMORY
ALTERA AVALON ONCHIP MEMORY
0x | 0100 | 0800 | boot ROM end | ALTERA AVALON ONCHIP MEMORY
0x | 0110 | 0000 | UART1 base address | ALTERA AVALON UART1 RS232
0x | 0110 | 0020 | UART USB base address | ALTERA AVALON UART2 USB
0x | 0120 | 0000 | Timer1 base address | ALTERA AVALON TIMER
0x | 0120 | 0020 | Timer2 base address | ALTERA AVALON TIMER
0x | 0120 | 0040 | Timer3 base address | ALTERA AVALON TIMER
0x | 0130 | 0000 | flash ry by base address | ALTERA AVALON PIO
0x | 0130 | 0010 | test in | ALTERA AVALON PIO
0x | 0130 | 0020 | test out | ALTERA AVALON PIO
0x | 0130 | 0030 | DMA enable | ALTERA AVALON PIO
0x | 0130 | 0040 | serial data adr. | ALTERA AVALON PIO
0x | 0130 | 0050 | serial data sw. | ALTERA AVALON PIO
0x | 0130 | 0060 | serial LED | ALTERA AVALON PIO
0x | 0130 | 0070 | ser. DAC base address | ALTERA AVALON PIO
0x | 0130 | 0080 | rotary interrupt base address | ALTERA AVALON PIO
0x | 0130 | 0090 | rotary interrupt ready base address | ALTERA AVALON PIO
0x | 0130 | 00A0 | speed read base address | ALTERA AVALON PIO
0x | 0130 | 00B0 | key base address | ALTERA AVALON PIO
0x | 0130 | 00C0 | key interrupt | ALTERA AVALON PIO
0x | 0130 | 00D0 | reset key interrupt | ALTERA AVALON PIO
0x | 0130 | 00E0 | PWM DAC base address | ALTERA AVALON PIO
0x | 0130 | 00F0 | mode | ALTERA AVALON PIO
0x | 0130 | 0100 | data adr. | ALTERA AVALON PIO
0x | 0130 | 0110 | start acquisition base address | ALTERA AVALON PIO
0x | 0130 | 0120 | data acquisition interrupt | ALTERA AVALON PIO
0x | 0130 | 0130 | button PIO | ALTERA AVALON PIO
0x | 0130 | 0140 | LA data PIO | ALTERA AVALON PIO
0x | 0130 | 0150 | LA interrupt | ALTERA AVALON PIO
0x | 0130 | 0160 | LA gate out | ALTERA AVALON PIO
0x | 0130 | 0170 | LA gate pulse | ALTERA AVALON PIO
0x | 0130 | 0180 | read vsync | ALTERA AVALON PIO
0x | 0130 | 0190 | boot reset key | ALTERA AVALON PIO
0x | 0130 | 01F0 |  | ALTERA AVALON PIO
0x | 0130 | 0200 | power on boot base address | ALTERA AVALON PIO

## Program Structure
Subroutine level 0 | Subroutine level 1 | Subroutine level 2 | Subroutine level 3 | Subroutine level 4
Program Start
Hardware::Init()
Display::Init()
Hardware::Restore_From_Flash()
Hardware::Start_Up()
Main Loop Start
Hardware::Rotary_Interface();
Display::REMOVE_SPLASH( )
Display::DrawStartScreen( )
Display::StatusUpdate( )
Display::DRAWCURSORDATA(0)
Display::DRAWQMDATA(0)
Display::DRAW_ALL_CURSOR( )
Start_Record( )
ResetTimer( )
ON_Rotary_Position( )
Recalc_Vars( )
Display::CALCPRETRIGGER( )
Display::DRAW_ALL_CURSOR()
ON_Trigger_Level()
UpdateTrigger(17)
Display::DRAWSTATUS(6, 1)
ON_Main_Wheel()
HandleMainWheel( )
MenuButtonCheck(3)
Display::DRAWP( )
Display::TEXTOUTxvbig( )
PIXELP( )
Display::DRAWSTATUS( )
Display::CALCCURSORDATA( )
sprintf( )
RenderText( )
Display::CALCQMDATA( )
Display::StatusUpdate( )
Recalc_Trigger( )
SetupADC( )
Display::DRAWMEMORY( )
Display::CALCPRETRIGGER( )
Display::DRAW_ALL_CURSOR( )
FloatStr::RenderText( )
ON_Timebase()
MenuButtonCheck(-1)
Recalc_Vars( )
Recalc_Trigger( )
Display::CALCCURSORDATA( )
Display::DRAWMEMORY( )
Display::CALCPRETRIGGER( )
Display::CALCCURSORDATA( )
Display::DRAW_ALL_CURSOR( )
ON_Zero_Channel_x( ) x= 1,2,3,4
ON_Voltage_Channel_x( ) x= 1,2,3,4
Hardware::Keyboard_Interface();
ResetTimer( )
Buttonhandler( )
Display::REMOVE_SPLASH()
Display::DrawStartScreen()
Display::StatusUpdate()
Display::DRAWCURSORDATA(0)
Display::DRAWQMDATA(0)
Display::DRAW_ALL_CURSOR( )
Start_Record( )
Display::RemoveStatusBtn( )
Display::DRAWMENUPOPDOWN( )
ClearPlanes( )
Display::CALCCURSORDATA( )
Display::StatusUpdate( )
MenuButtonCheck( )
Set_LED( )
Reset_LED( )
Send_LED( )
UpdateTrigger( )
Display::CALCQMDATA( )
SendData( )
AutoScale( )
F1_Btnhandler( )
Display::DRAWMENUPOPDOWN( )
Display::DRAWMENUPOPUP( )
SetSwitches( )
Hardware::SearchZeros( )
ON_Zero_Channel_n( ) n = 1,2
Display::GRID( )
SetupADC( )
Display::StatusUpdate( )
ClearPlanes( )
Display::DRAW_ALL_CURSOR( )
UpdateTrigger( )
SendData( )
Display::CALCCURSORDATA( )
Display::CALCQMDATA( )
F2_Btnhandler( )
F3_Btnhandler( )
F4_Btnhandler( )
F5_Btnhandler( )
F6_Btnhandler( )
Start_Record( )
-> Test functions
SendData( )
Hardware::Handle_ADC();
READADC( )
WRITEADC( )
ReadOut_Signal( )
READADC( )
READADC_ALL2( )
PREPARE_READADC( )
READADC_ALL( )
Display::ProcessSignalData( );
FFT_CalcSpectrum( )
MathCalculation
Roll/ShiftMode
NoiseFilter
Interpolate( )
Display::DRAWSIGNALS(1)
DRAWSIGNAL_XY( )
PIXELP( )
DRAWLINEP( )
DRAWSIGNAL_TY( )
PIXELP( )
Connect_Pixels( )
DRAWLINEP( )
DRAWSIGNAL_FY( )
Connect_Pixels( )
DRAWZEROLINE( )
CALCQMDATA( )
FloatStr::RenderText( )
Hardware::TransferPlanes();
TransferPlanes_asm_persistant( )
TransferPlanes_asm( )
TransferPlane_Set_Buffer_Adr( )
Display::DRAWSIGNS( )
RemoveZeroSign( )
CLEAR2BYTE( )
RemoveZeroSignBack( )
CLEAR2BYTEBACK( )
DrawZeroSign( )
WRITE2BYTE( )
DrawZeroSignBack( )
WRITE2BYTEBACK( )
Display::UpdateMenu( )
ClearMenuItemAsm( )
DrawMenuItemButton( )
ClearMenuItemTextAsm( )
DrawMenuItem( )
DRAWCURSORDATA( )
Hardware::Transfer_USB_Data();
AMDFlash::Write_Config_Flash();

## LED Table
LED No. | Function | Description
0 | Channel 1 | Yellow
1 | Channel 2 | Green
2 | Channel 3 | Blue
3 | Channel 4 | Red
4 | Math | White
5 | Main Wheel | White
6 | Cursor | White
7 | Quick Measure | White
8 | Edge | White
9 | Trigger armed | Trigger is armed (green)
10 | Pulse Width | White
11 | Trigger event | Trigger event occured (red)
12 | Run/Stop | Green
13 | Run/Stop | Red
14 | Single | Red
15 | Single | Green

## Button Table
Btn No. | Function | Description
0 | Acquire
1 | Channel 1
2 | Channel 2
3 | Channel 3
4 | Channel 4
5 | Display
6 | Utility
7
8 | Save/Recall
9
10 | Main/Delayed | Calls mode menu (main/delayed/XY)
11
12 | Mode/Coupling | Signal coupling and mode selection (Auto/Normal)
13
14 | Edge | Trigger menu
15 | Pulse Width
16 | Cursor
17 | Quick Print
18
19 | Quick Measure
20
21
22
23 | Auto-scale
24
25
26 | Math
27
28
29
30 | Function 1 | UserIF::F1_Btnhandler()
31 | Function 2 | UserIF::F2_Btnhandler()
32 | Function 3 | UserIF::F3_Btnhandler()
33 | Function 4 | UserIF::F4_Btnhandler()
34 | Function 5 | UserIF::F5_Btnhandler()
35 | Function 6 | UserIF::F6_Btnhandler()
36
37 | Run/Stop
38 | Single
39

## Menu Table
Menu
No. | Function | F1 [0] | F2 [1] | F3 [2] | F4 [3] | F5 [4] | F6 [5] | Description
0 | Acquire | ^ Logic | ^ Averaging | ^ Noise Filter | ^ Interpolation |  |  | Acquire menu,
1 | Channel 1 | ^ Coupling | BW limit | Invert | Probe | Center Channel 1 | Dispatch Channels | Channel dependent settings
2 | Channel 2 | ^ Coupling | BW limit | Invert | Probe | Center Channel 2 | Dispatch Channels | Channel dependent settings
3 | Channel 3 | ^ Coupling | BW limit | Invert | Probe | Center Channel 3 | Dispatch Channels | Channel dependent settings
4 | Channel 4 | ^ Coupling | BW limit | Invert | Probe | Center Channel 4 | Dispatch Channels | Channel dependent settings
5 | Display | Persistent | Clear Display | Grid Intensity | ^ No Display | ^ Draw Mode | Setup ↓ | Display settings in time and X-Y mode
6 | Utility | About Oscilloscope | Default Setup | Calibrate Offsets | ^ Calibration Set | Test Signal | More ↓ | Adjustment and test
7 | Service | Start User Cal | Start Self Test | User Cal Status | RS-232 Setup |  |  | not available
8 | Save/Recall | Memory | Save Trace | Recall Trace | Overlay Trace | Clear Display | Restore Settings | Save and load signals
9 | Display Setup | ^ Grid Color | ^ Grid Layout | ^ Status Layout | QM Cursors |  | Back ↑ | Display Setup → color palettes etc.
10 | Timebase | Main | Delayed | X-Y | FFT | Browse | USTB↓ | Timebase menu
11 | Delayed Timebase | [6 empty] | not available
12 | Trigger Mode | ^ Mode | ^ Coupling | Auto Level | Holdoff | External Probe | More ↓ | Trigger settings 1 → Mode/Coupling
13 | Trigger Sub | ^ LED1 | ^ LED2 | Map to Ch 3+4 | ^ Man.Trig. Setup. | ^ Pre Trigger Setup | Back ↑ | Trigger settings submenu
14 | Trigger Edge | Rising / Falling | ^ Trigger Source | External | TV | Adjust Window | PreTrigger | Trigger settings 2 → Edge triggering
15 | Pulse Width | ^ Source | Pulse | Direction | Width | Width | PreTrigger | Trigger settings 3 → Pulse Width triggering
16 | Cursor | ^ Source | X1 | X2 | Y1 | Y2 | X1X2/Y1Y2 | Cursor settings
17 | Quick Print | Save to BMP | Save to PPM | Save to CSV | Save to ASCII | Send Measure | ^ Destination | Data export and screenshots on PC
18 | Hardware | ^ADC Setup | ^Pre Amp Gain | ^ADC Driver | LED Test |  | Back ↑ | Special hardware settings submenu
19 | Quick Measurement | ^ Source | Select Measurement | Measure | Clear Measure | Setting ↓ | Thresholds ↓ | Quick measure settings
20 | QM Settings Delay | Source | Edge | Source | Edge | Measure Delay | Back ↑ | QM → special option delay measurement
21 | QM Settings Phase | Source |  | Source |  | Measure Phase | Back ↑ | QM → special option phase measurement
22 | QM Thresholds | ^ Source | Type | Lower | Middle | Upper | Back ↑ | More QM-settings
23 | Autoscale | Undo | Search slow Timebase | Save new Settings | ^ TB Search |  |  | Autoscale functions
24 | Trigger Mode TV | Mode | Pos/Neg | Standard | Mode | Count |  | not available
25 | Extended Utilities | CH1 Delay | CH2 Delay | CH3 Delay | CH4 Delay | Hardware ↓ | Back ↑ | Channel delay and hardware submenu
26 | Math | Scale | Offset | 1 * 2 | 1 - 2 | 1 + 2 |  | Mathematic functions
27 | FFT | ^ Source | ^ Window | ^ Mode | ^ Length |  | Back ↑ | FFT settings submenu
28 | Display FFT | Persistent | Clear Display | Grid Intensity | Switch Grid |  | Setup ↓ | FFT display settings
29 | [7 empty] | not used
30 | USTB | ^ USTB Mode | ^ USTB Display | ^ Buffer Size |  |  | Back ↑ | Submenu USTB settings
31 | Quick Print USB | Save to color | Save to B/W | Save to CSV |  |  | ^ Destination | Data export to SD card / USB-host

## Timbase Table
Timebase | Ta | Factor | Sample Depth | Memory | Main TB | TB-Register | Delayed Timebase | Zoom-Factor | Length
2ns | 1GSa/s | 1 | 16384 | 16KB | 0 | 0xFFFFFFFF | n.a. | 0,04 | 24
5ns | 1GSa/s | 1 | 16384 | 16KB | 1 | 0xFFFFFFFF | 2ns | 0,1 | 60
10ns | 1GSa/s | 1 | 16384 | 16KB | 2 | 0xFFFFFFFF | 5ns /2ns | 0,2 | 120
20ns | 1GSa/s | 1 | 16384 | 16KB | 3 | 0xFFFFFFFF | 10ns /5ns /2ns | 0,4 | 240
50ns | 1GSa/s | 1 | 16384 | 16KB | 4 | 0xFFFFFFFF | 20ns / 10ns /5ns | 1,0 | 600
100ns | 1GSa/s | 2 | 8192 | 16KB | 5 | 0xFFFFFFFF | 50ns / 20ns /10ns | 2,0 | 1200
200ns | 1GSa/s | 4 | 4096 | 16KB | 6 | 0xFFFFFFFF | 100ns / 50ns / 20ns /10ns | 4,0 | 2400
500ns | 1GSa/s | 10 | 1638 | 16KB | 7 | 0xFFFFFFFF | 200ns / 100ns / 50ns / 20ns /10ns /5ns | 10,0 | 6000
1µs | 1GSa/s | 20 | 819 | 16KB | 8 | 0xFFFFFFFF | 500ns / 200ns / 100ns / 50ns / 20ns /10ns | 20,0 | 12000
2µs | 500MSa/s | 20 | 819 | 16KB | 9 | 0xFFFFFFFE | 1µs / 500ns / 200ns / 100ns | 20,0 | 12000
5µs | 250MSa/s | 25 | 655 | 16KB | 10 | 0xFFFFFFFD | 2µs / 1µs / 500ns / 200ns | 25,0 | 15000
10µs | 25MSa/s | 5 | 819 | 4KB | 11 | 0xFFFFFFFB | 5µs / 2µs | 5,0 | 3000
20µs | 10MSa/s | 4 | 1024 | 4KB | 12 | 0xFFFFFFF3 | 10µs / 5µs | 4,0 | 2400
50µs | 5MSa/s | 5 | 819 | 4KB | 13 | 0xFFFFFFE7 | 20µs / 10µs | 5,0 | 3000
100µs | 2,5MSa/s | 5 | 819 | 4KB | 14 | 0xFFFFFFCE | 50µs / 20µs | 5,0 | 3000
200µs | 1MSa/s | 4 | 1024 | 4KB | 15 | 0xFFFFFF83 | 100µs / 50µs | 4,0 | 2400
500µs | 500KSa/s | 5 | 819 | 4KB | 16 | 0xFFFFFF06 | 200µs / 100µs | 5,0 | 3000
1ms | 250KSa/s | 5 | 819 | 4KB | 17 | 0xFFFFFE0C | 500µs / 200µs | 5,0 | 3000
2ms | 100KSa/s | 4 | 1024 | 4KB | 18 | 0xFFFFFB1E | 1ms / 500µs | 4,0 | 2400
5ms | 50KSa/s | 5 | 819 | 4KB | 19 | 0xFFFFF63C | 2ms / 1ms | 5,0 | 3000
10ms | 25KSa/s | 5 | 819 | 4KB | 20 | 0xFFFFEC78 | 5ms / 2ms | 5,0 | 3000
20ms | 10KSa/s | 4 | 1024 | 4KB | 21 | 0xFFFFCF2C | 10ms / 5ms | 4,0 | 2400
50ms | 5KSa/s | 5 | 819 | 4KB | 22 | 0xFFFF9E58 | 20ms / 10ms | 5,0 | 3000
100ms | 2,5KSa/s | 5 | 819 | 4KB | 23 | 0xFFFF3CB0 | 50ms / 20ms | 5,0 | 3000
200ms | 1KSa/s | 4 | 1024 | 4KB | 24 | 0xFFFE17B8 | 100ms / 50ms | 4,0 | 2400
500ms | 500Sa/s | 5 | 819 | 4KB | 25 | 0xFFFC2F70 | 200ms / 100ms | 5,0 | 3000
1s | 50Sa/s | 1 | roll/shift mode | 32KB /16KB | 26 | 0xFFFFFFFF | n.a. | 1,0 | 600
2s | 25Sa/s | 1 | roll/shift mode | 32KB /16KB | 27 | 0xFFFFFFFF | n.a. | 1,0 | 600
5s | 10Sa/s | 1 | roll/shift mode | 32KB /16KB | 28 | 0xFFFFFFFF | n.a. | 1,0 | 600
10s | 5Sa/s | 1 | roll/shift mode | 32KB /16KB | 29 | 0xFFFFFFFF | n.a. | 1,0 | 600
20s | 2.5Sa/s | 1 | roll/shift mode | 32KB /16KB | 30 | 0xFFFFFFFF | n.a. | 1,0 | 600
50s | 1Sa/s | 1 | roll/shift mode | 32KB /16KB | 31 | 0xFFFFFFFF | n.a. | 1,0 | 600
100s | 1Sa/2s | 1 | roll/shift mode | 32KB /16KB | 32 | 0xFFFFFFFF | n.a. | 1,0 | 600
200s | 1Sa/4s | 1 | roll/shift mode | 32KB /16KB | 33 | 0xFFFFFFFF | n.a. | 1,0 | 600
500s | 1Sa/10s | 1 | roll/shift mode | 32KB /16KB | 34 | 0xFFFFFFFF | n.a. | 1,0 | 600
1000s | 1Sa/20s | 1 | roll/shift mode | 32KB /16KB | 35 | 0xFFFFFFFF | n.a. | 1,0 | 600
Ta - sampling rate
Factor - number of real samples that are dropped out between two shown samples (step width)
Sample Depth - number of used samples from the 16384 or 4096 real samples (Memory / Factor)
Compressed or interpolated timebase
Compressed or interpolated timebase
Real timebase

## Delayed Tb Table
Timebase | Delayed Tb | Ta | Factor | Del. Factor | Tb Ratio | Mein TB | Delayed TB
2ns | n.a. | 1GSa/s | 1(0,04) | n.a. | n.a. | 0 | 0
5ns | 2ns | 1GSa/s | 1(0,1) | 1(0,04) | 2,5 | 1 | 0
10ns | 2ns | 1GSa/s | 1(0,2) | 1(0,04) | 5 | 2 | 1
10ns | 5ns | 1GSa/s | 1(0,2) | 1(0,1) | 2 | 2 | 0
20ns | 2ns | 1GSa/s | 1(0,4) | 1(0,04) | 10 | 3 | 2
20ns | 5ns | 1GSa/s | 1(0,4) | 1(0,1) | 4 | 3 | 1
20ns | 10ns | 1GSa/s | 1(0,4) | 1(0.2) | 2 | 3 | 0
50ns | 5ns | 1GSa/s | 1 | 1(0,1) | 10 | 4 | 2
50ns | 10ns | 1GSa/s | 1 | 1(0,2) | 5 | 4 | 1
50ns | 20ns | 1GSa/s | 1 | 1(0,4) | 2,5 | 4 | 0
100ns | 5ns | 1GSa/s | 2 | 1(0,1) | 20 | 5 | 3
100ns | 10ns | 1GSa/s | 2 | 1(0,2) | 10 | 5 | 2
100ns | 20ns | 1GSa/s | 2 | 1(0,4) | 5 | 5 | 1
100ns | 50ns | 1GSa/s | 2 | 1 | 2 | 5 | 0
200ns | 5ns | 1GSa/s | 4 | 1(0,1) | 40 | 6 | 4
200ns | 10ns | 1GSa/s | 4 | 1(0,2) | 20 | 6 | 3
200ns | 20ns | 1GSa/s | 4 | 1(0,4) | 10 | 6 | 2
200ns | 50ns | 1GSa/s | 4 | 1 | 4 | 6 | 1
200ns | 100ns | 1GSa/s | 4 | 2 | 2 | 6 | 0
500ns | 5ns | 1GSa/s | 10 | 1(0,1) | 100 | 7 | 5
500ns | 10ns | 1GSa/s | 10 | 1(0,2) | 50 | 7 | 4
500ns | 20ns | 1GSa/s | 10 | 1(0,4) | 25 | 7 | 3
500ns | 50ns | 1GSa/s | 10 | 1 | 10 | 7 | 2
500ns | 100ns | 1GSa/s | 10 | 2 | 5 | 7 | 1
500ns | 200ns | 1GSa/s | 10 | 4 | 2,5 | 7 | 0
1µs | 10ns | 1GSa/s | 20 | 1(0,2) | 100 | 8 | 5
1µs | 20ns | 1GSa/s | 20 | 1(0,4) | 50 | 8 | 4
1µs | 50ns | 1GSa/s | 20 | 1 | 20 | 8 | 3
1µs | 100ns | 1GSa/s | 20 | 2 | 10 | 8 | 2
1µs | 200ns | 1GSa/s | 20 | 4 | 5 | 8 | 1
1µs | 500ns | 1GSa/s | 20 | 10 | 2 | 8 | 0
2µs | 100ns | 500MSa/s | 20 | 1 | 20 | 9 | 3
2µs | 200ns | 500MSa/s | 20 | 2 | 10 | 9 | 2
2µs | 500ns | 500MSa/s | 20 | 5 | 4 | 9 | 1
2µs | 1µs | 500MSa/s | 20 | 10 | 2 | 9 | 0
5µs | 200ns | 250MSa/s | 25 | 1 | 25 | 10 | 3
5µs | 500ns | 250MSa/s | 25 | 2,5 | 10 | 10 | 2
5µs | 1µs | 250MSa/s | 25 | 5 | 5 | 10 | 1
5µs | 2µs | 250MSa/s | 25 | 10 | 2,5 | 10 | 0
10µs | 2µs | 25MSa/s | 5 | 1 | 5 | 11 | 1
10µs | 5µs | 25MSa/s | 5 | 2,5 | 2 | 11 | 0
20µs | 5µs | 10MSa/s | 4 | 1 | 4 | 12 | 1
20µs | 10µs | 10MSa/s | 4 | 2 | 2 | 12 | 0
50µs | 10µs | 5MSa/s | 5 | 1 | 5 | 13 | 1
50µs | 20µs | 5MSa/s | 5 | 2 | 2,5 | 13 | 0
100µs | 20µs | 2,5MSa/s | 5 | 1 | 5 | 14 | 1
100µs | 50µs | 2,5MSa/s | 5 | 2,5 | 2 | 14 | 0
200µs | 50µs | 1MSa/s | 4 | 1 | 4 | 15 | 1
200µs | 100µs | 1MSa/s | 4 | 2 | 2 | 15 | 0
500µs | 100µs | 500KSa/s | 5 | 1 | 5 | 16 | 1
500µs | 200µs | 500KSa/s | 5 | 2 | 2,5 | 16 | 0
1ms | 200µs | 250KSa/s | 5 | 1 | 5 | 17 | 1
1ms | 500µs | 250KSa/s | 5 | 2,5 | 2 | 17 | 0
2ms | 500µs | 100KSa/s | 4 | 1 | 4 | 18 | 1
2ms | 1ms | 100KSa/s | 4 | 2 | 2 | 18 | 0
5ms | 1ms | 50KSa/s | 5 | 1 | 5 | 19 | 1
5ms | 2ms | 50KSa/s | 5 | 2 | 2,5 | 19 | 0
10ms | 2ms | 25KSa/s | 5 | 1 | 5 | 20 | 1
10ms | 5ms | 25KSa/s | 5 | 2,5 | 2 | 20 | 0
20ms | 5ms | 10KSa/s | 4 | 1 | 4 | 21 | 1
20ms | 10ms | 10KSa/s | 4 | 2 | 2 | 21 | 0
50ms | 10ms | 5KSa/s | 5 | 1 | 5 | 22 | 1
50ms | 20ms | 5KSa/s | 5 | 2 | 2,5 | 22 | 0
100ms | 20ms | 2,5KSa/s | 5 | 1 | 5 | 23 | 1
100ms | 50ms | 2,5KSa/s | 5 | 2,5 | 2 | 23 | 0
200ms | 50ms | 1KSa/s | 4 | 1 | 4 | 24 | 1
200ms | 100ms | 1KSa/s | 4 | 2 | 2 | 24 | 0
500ms | 100ms | 500Sa/s | 5 | 1 | 5 | 25 | 1
500ms | 200ms | 500Sa/s | 5 | 2 | 2,5 | 25 | 0

## WELEC USB Interface
WELEC USB Transmission Protocol Main Functions
RX byte 0 | RX byte 1 | Function
block size 0x25 (37) | 0x00 (0) | noop
block size 0x25 (37) | 0x01 (1) | send screenshot
block size 0x25 (37) | 0x02 (2) | send data | Function 1 - 22 | not used
block size 0x25 (37) | 0x03 (3) | receive data | Function 0 - 23 | data
block size 0x25 (37) | 0x04 (4) | not used
block size 0x25 (37) | 0x05 (5) | start record
block size 0x25 (37) | 0x06 (6) | receive commands (subfunctions 1 - 35 in RX byte 2) | command 1 - 35
block size 0x25 (37) | 0x07 (7) | test functions to set the LEDs on the front panel
block size 0x25 (37) | 0x08 (8) | receive commands (subfunctions 1 - 12 in RX byte 2)
block size 0x25 (37) | 0x09 (9) | no function
block size 0x25 (37) | 0x0A (10) | read and write the flash memory (subfunctions 0 - 18 in RX byte 2)
block size 0x25 (37) | 0x0B (11) | set timebase register (test function)
block size 0x25 (37) | 0x0C (12) | send timebase register to PC (test function)
block size 0x25 (37) | 0x0D (13) | search zeroes
block size 0x25 (37) | 0x0E (14) | send all data to PC (refresh)
block size 0x25 (37) | 0x0F (15) | set DSO to default parameters
block size 0x25 (37) | 0x10 (16) | send all DSO parameters to PC (connect)
block size 0x25 (37) | 0x11 (17) | send display parameters to PC
block size 0x25 (37) | 0x12 (18) | no function
WELEC W2000PC USB Commands
Command | RX byte 0 | RX byte 1 | RX byte 2 | RX byte 3 | RX byte 4 - 36
connect to DSO | 0x25 (37) | 0x10 (16) | no function | no function | no function
channel 1 on/off | 0x25 (37) | 0x06 (6) | 0x07 (7) | 0x01 (1) / 0x00 (0) | no function
channel 2 on/off | 0x25 (37) | 0x06 (6) | 0x09 (9) | 0x01 (1) / 0x00 (0) | no function
channel 3 on/off | 0x25 (37) | 0x06 (6) | 0x0A (10) | 0x01 (1) / 0x00 (0) | no function
channel 4 on/off | 0x25 (37) | 0x06 (6) | 0x0B (11) | 0x01 (1) / 0x00 (0) | no function
stop | 0x25 (37) | 0x06 (6) | 0x18 (24) | 0x01 (1) | no function
single shot | 0x25 (37) | 0x06 (6) | 0x19 (25) | 0x01 (1) | no function
channel 1 AC/DC | 0x25 (37) | 0x06 (6) | 0x2C (28) | 0x01 (1) / 0x00 (0) | no function
channel 2 AC/DC | 0x25 (37) | 0x06 (6) | 0x2D (29) | 0x01 (1) / 0x00 (0) | no function
channel 3 AC/DC | 0x25 (37) | 0x06 (6) | 0x2E (30) | 0x01 (1) / 0x00 (0) | no function
channel 4 AC/DC | 0x25 (37) | 0x06 (6) | 0x2F (31) | 0x01 (1) / 0x00 (0) | no function
Autoscale | 0x25 (37) | 0x06 (6) | 0x0C (12) | no function | no function
timebase wheel right/left | 0x25 (37) | 0x08 (8) | 0x0B (11) | 0x01 (1) / 0x02 (2) | no function
channel 1 voltage up/down | 0x25 (37) | 0x08 (8) | 0x05 (5) | 0x01 (1) / 0x02 (2) | no function
channel 2 voltage up/down | 0x25 (37) | 0x08 (8) | 0x06 (6) | 0x01 (1) / 0x02 (2) | no function
channel 3 voltage up/down | 0x25 (37) | 0x08 (8) | 0x07 (7) | 0x01 (1) / 0x02 (2) | no function
channel 4 voltage up/down | 0x25 (37) | 0x08 (8) | 0x08 (8) | 0x01 (1) / 0x02 (2) | no function
send all data to PC (refresh) | 0x25 (37) | 0x0E(14) | no function | no function | no function
WELEC W2000-Update USB Commands
Command | RX byte 0 | RX byte 1 | RX byte 2 | RX byte 3 | RX byte 4 - 36
disable/enable programming | 0x25 (37) | 0x0A (10) | 0x00 (0) | 0x00 (0) / 0x01 (1) | no function
erase flash sector | 0x25 (37) | 0x0A (10) | 0x01 (1) | sector address | sector address
flash adress offset | 0x25 (37) | 0x0A (10) | 0x02 (2) | address offset | address offset
write data block to flash | 0x25 (37) | 0x0A (10) | 0x03 (3) | data | data, block size
read data from flash (test) | 0x25 (37) | 0x0A (10) | 0x04 (4) | address | address
send firmware version to PC | 0x25 (37) | 0x0A (10) | 0x05 (5) | no function | no function
send hardware version to PC | 0x25 (37) | 0x0A (10) | 0x06 (6) | no function | no function
read data from flash (test) | 0x25 (37) | 0x0A (10) | 0x07 (7) | address | address
read data from flash (test) | 0x25 (37) | 0x0A (10) | 0x04 (8) | address | address
read data from flash (test) | 0x25 (37) | 0x0A (10) | 0x04 (9) | address | address
read data from flash (test) | 0x25 (37) | 0x0A (10) | 0x04 (9) | address | address
read data from flash (test) | 0x25 (37) | 0x0A (10) | 0x04 (10) | address | address
read data from flash (test) | 0x25 (37) | 0x0A (10) | 0x04 (11) | address | address
read data from flash (test) | 0x25 (37) | 0x0A (10) | 0x04 (12) | address | address
read data from flash (test) | 0x25 (37) | 0x0A (10) | 0x04 (13) | address | address
read data from flash (test) | 0x25 (37) | 0x0A (10) | 0x04 (14) | address | address
read data from flash (test) | 0x25 (37) | 0x0A (10) | 0x04 (15) | address | address
read data from flash (test) | 0x25 (37) | 0x0A (10) | 0x04 (16) | address | address
read data from flash (test) | 0x25 (37) | 0x0A (10) | 0x04 (17) | address | address
read data from flash (test) | 0x25 (37) | 0x0A (10) | 0x00 (18) | address | address

## BF USB Interface
Blue Flash USB Transmission Protocol Main Functions
RX byte 0 (block size) | RX byte 1 | Function
0x08 (8) | 0x00 (0) | noop
0x08 (8) | 0x01 (1) | command mode (remote control)
0x08 (8) / 0x40 (64) | 0x02 (2) | write to flash (firmware upload, set protected config etc.) | Function 1 - 22 | not used
0x40 (64) | 0x03 (3) | read from flash | Function 0 - 23 | data
0x08 (8) | 0x04 (4) | test function
Remote Control USB Commands
Command | RX byte 0 | RX byte 1 | RX byte 2 | RX byte 3 | RX byte 4 - 7
USB connection request | 0x08 (8) | 0x01 (1) | 0x00 (0) | 0x55 (85) | no function
channel 1 button | 0x08 (8) | 0x01 (1) | 0x01 (1) | no function | no function
channel 2 button | 0x08 (8) | 0x01 (1) | 0x02 (2) | no function | no function
channel 3 button | 0x08 (8) | 0x01 (1) | 0x03 (3) | no function | no function
channel 4 button | 0x08 (8) | 0x01 (1) | 0x04 (4) | no function | no function
math button | 0x08 (8) | 0x01 (1) | 0x05 (5) | no function | no function
| 0x08 (8) | 0x01 (1) | 0x06
| 0x08 (8) | 0x01 (1) | 0x07
| 0x08 (8) | 0x01 (1) | 0x08
| 0x08 (8) | 0x01 (1) | 0x09
| 0x08 (8) | 0x01 (1) | 0x10
| 0x08 (8) | 0x01 (1) | 0x11
| 0x08 (8) | 0x01 (1) | 0x12
| 0x08 (8) | 0x01 (1) | 0x13
| 0x08 (8) | 0x01 (1) | 0x14
| 0x08 (8) | 0x01 (1) | 0x15
| 0x08 (8) | 0x01 (1) | 0x16
| 0x08 (8) | 0x01 (1) | 0x17
| 0x08 (8) | 0x01 (1) | 0x18
| 0x08 (8) | 0x01 (1) | 0x19
| 0x08 (8) | 0x01 (1) | 0x20
function button 1 | 0x08 (8) | 0x01 (1) | 0x21 (33) | no function | no function
function button 2 | 0x08 (8) | 0x01 (1) | 0x22 (34) | no function | no function
function button 3 | 0x08 (8) | 0x01 (1) | 0x23 (35) | no function | no function
function button 4 | 0x08 (8) | 0x01 (1) | 0x24 (36) | no function | no function
function button 5 | 0x08 (8) | 0x01 (1) | 0x25 (37) | no function | no function
function button 6 | 0x08 (8) | 0x01 (1) | 0x26 (38) | no function | no function
| 0x08 (8) | 0x01 (1) | 0x27
| 0x08 (8) | 0x01 (1) | 0x28
| 0x08 (8) | 0x01 (1) | 0x29
| 0x08 (8) | 0x01 (1) | 0x30
| 0x08 (8) | 0x01 (1) | 0x31
| 0x08 (8) | 0x01 (1) | 0x32
| 0x08 (8) | 0x01 (1) | 0x33
| 0x08 (8) | 0x01 (1) | 0x34
| 0x08 (8) | 0x01 (1) | 0x35
| 0x08 (8) | 0x01 (1) | 0x36
| 0x08 (8) | 0x01 (1) | 0x37
Flash Writing USB Commands
Command | RX byte 0 | RX byte 1 | RX byte 2 | RX byte 3 | RX byte 4 - 63
disable/enable programming | 0x08 (8) | 0x02 (2) | 0x00 (0) | 0x00 (0) / 0x01 (1) | no function
erase flash sector | 0x08 (8) | 0x02 (2) | 0x01 (1) | sector address | sector address
write data block to flash | 0x40 (64) | 0x02 (2) | 0x03 (3) | address | address / data
Flash Reading USB Commands
Command | RX byte 0 | RX byte 1 | RX byte 2 | RX byte 3 | RX byte 4 - 63
| 0x40 (64) | 0x02 (2) | 0x00 (0)
| 0x40 (64) | 0x02 (2) | 0x01 (1)
read data block from flash | 0x40 (64) | 0x02 (2) | 0x03 (3) | address | address / data

## FFT Ranges
Range | A max | A min | dBm max | dBm min
100V | 400V | 2V | 62,04dBm
50V | 200V | 1V | 56,02dBm
20V | 80V | 400mV | 48,06dBm
10V | 40V | 200mV
5V | 20V | 100mV
2V | 8V | 40mV
1V | 4V | 20mV
500mV | 2V | 10mV
200mV | 800mV | 4mV
100mV | 400mV | 2mV
50mV | 200mV | 1mV
20mV | 80mV | 400µV
10mV | 40mV | 200µV
5mV | 20mV | 100µV
2mV | 8mV | 40µV
1mV | 4mV | 20µV