RIGOL DSA815-TG - 2014-09-12

Content

Hidden Modes - Key and ScanCode

SCPI commands

SCPI commands list

SCPI parameter list

Non-visible menu items

SDRAM

Hidden Modes - Graphic

Correction file

Hidden Modes - Keys and ScanCode

Factory Boot Mode

BW          52 00 00 00

SWEEP       62 00 00 00

MARK        12 01 00 00

TG          e2 00 00 00

MEAS        f2 00 00 00

MEASSET     02 01 00 00

DEMOD       c2 00 00 00

Requires an inserted USB flash drive.

Probably requires an appropriate firmware on the USB stick.

Maintenance Mode

TRACE       d2 00 00 00

TG          e2 00 00 00

MARKFUNC    22 01 00 00

MEASSET     02 01 00 00

SYSTEM      82 01 00 00

PRINTSETUP  c2 01 00 00

STORAGE     92 01 00 00

Adds menu: System -> Service - Do not use Calibration -> Clr Int Cal

Adds function: TG -> ALC

Unknown Mode 1

BW          52 00 00 00

MARK        12 01 00 00

TRACE       d2 00 00 00

TG          e2 00 00 00

MARKFUNC    22 01 00 00

SWEEP       62 00 00 00

TG          e2 00 00 00

Unknown Mode 2

MEASSET     02 01 00 00

PRINTSETUP  c2 01 00 00

TG          e2 00 00 00

MEAS        f2 00 00 00

SYSTEM      82 01 00 00

TG          e2 00 00 00

DEMOD       c2 00 00 00

If the keys for „Unknown Mode 1“, „Unknown Mode 2“ and STORAGE is entered in the menu System -> Self-Test -> Key Test, a chinese text is displayed:

DSA系列产品声明

本产品版权为普源精电科技有限公司所有，盗版必究!

Google translation:

DSA Series Product Declaration

This product is copyright Rigol Technologies Ltd. All pirated reserved!

User Mode

STORAGE     92 01 00 00

ENTER       e2 05 00 00

ESC         d2 05 00 00

BACK        c2 05 00 00

FREQ        12 00 00 00

FREQ        12 00 00 00

FREQ        12 00 00 00

Exits the Maintenance Mode.

SCPI commands

Firmware v01.09

:PRIVate:FACTory:SN? -> "DSA8A145208137"

:PRIVate:FACTory:MODEL? -> "DSA815"

:PRIVate:FACTory:MAC? -> "00-19-AF-52-06-F0"

:PRIVate:FACTory:RTC? -> "2014-07-20 16:19:14"

:PRIVate:FACTory:CALRtc? -> "2012-08-14 06-48-56"

:PRIVate:FACTory:AKEY? x -> "0"

:PRIVate:FACTory:SKEY? -> Output only if set. Do not use, all options will be lost.

:PRIVate:CTRL:DATAtype? -> "0"

:PRIVate:CTRL:ESR? -> "1"

:PRIVate:CTRL:SYNC? -> (:SYST:ERR? -220, "Parameter error")

:PRIVate:CTRL:TRACe? -> 9026 Bytes Testtrace

:PRIVate:TEMPeratue:GET? -> "0.000000E+00"

:PRIV:SOFT:VER? x -> "00.01.09"

Menu System -> Service -> Calibration:

:PRIVate:CALibration:AMPLref? -> "1167"

:PRIVate:CALibration:FREQref? -> "1799"

:PRIVate:CALibration:LOdac? -> "0"

:PRIVate:CALibration:CLKdac? -> "0"

:PRIVate:CALibration:ONdac? -> "0"

:PRIVate:CALibration:OFFdac? -> "0"

Menu System -> Service -> Log Manage:

:SERVice:LOG:DEVIce? -> "PC", "UDISK", "RMT" (Output Device: PC, U-Disk, Rmt-Monitor)

:SERVice:LOG:CPState? -> "0"

:SELFtest:KEY -> Brings up the Local 'Self Test' Menu for Key Test.

:SELFtest:SYS -> No response, no error

:DBGCMD:STAT? TX -> "(u16HostFetchErr = 0),(u16FifoLenErr = 0),(u32NoDataCnt = 0),(u32WaitDataCnt = 0),(u16FmtDataErr = 0),(u16FmtTypeErr = 0),(u16StateMachineErr = 0),(u16DeadLockCnt = 0),(u16SysFmtErr = 0),(u16AddrAccessErr = 0),(u32TotalTxByte = 56),(u32CurTxByte = 56)"

:DBGCMD:STAT? RX -> "(u16InteruptCnt = 0),(u16WInQCnt = 0),(u16RxDataCnt = 0),(u16ProCnt = 3),(u16IBFullCnt = 0),"

:DBGCMD:STAT? Q -> "(u16NoFreeNode = 0),(u16QIsFull = 0),(u16QIsEmpty = 0),(u16RxBufOverflow = 0),(u16RxDataPtrNull = 0),(u16RxDataLenErr = 0),"

:DBGCMD:STAT? PARSE -> "(u16InvalidCharacter = 0),(u16KeywordNotFound = 2),(u16CmdNotFound = 0),(u16TooManyKeyword = 0),(u16LastKeywordOmitted = 0),(u16CmdSNErr = 0),(u16InvalidCallBackFunc = 0),(u16CmdExcuteErr = 0),"

:DBGCMD:STAT? CLR -> "Clear all dbg stat."

:DBGCMD:VER? -> "SCP Module Ver : 00.02.02.00"

:KEYBoard:KEYValue Keyname

Keyname:

AMP

BACK

BW

DEMOD

DIGIT0

DIGIT1

DIGIT2

DIGIT3

DIGIT4

DIGIT5

DIGIT6

DIGIT7

DIGIT8

DIGIT9

DOT

DOWN

ENTER

ESC

F1

F2

F3

F4

F5

F6

F7

FREQ

HELP

KNOBL

KNOBR

MARK

MARKFUNC

MARKTO

MEAS

MEASSET

PAGEDOWN

PEAK

PRESET

PRINT

PRINTSETUP

RETURN

SPAN

STORAGE

SWEEP

SYSTEM

TG

TRACE

TUNE

UP

For example, show the frequency menu:

:KEYBoard:KEYValue ESC;:KEYBoard:KEYValue FREQ

or enter Maintenance Mode:

:KEYB:KEYV ESC;:KEYB:KEYV TRACE;:KEYB:KEYV TG;:KEYB:KEYV MARKFUNC;:KEYB:KEYV MEASSET;:KEYB:KEYV SYSTEM;:KEYB:KEYV PRINTSETUP;:KEYB:KEYV STORAGE

:SYST:FREad? Address decimal, length decimal -> Read Flash memory

The DSA uses only 4 MB of the 8 MB.

Read the first 4 MB.:

:SYST:FRE? 0,1048576

:SYST:FRE? 1048576,1048576

:SYST:FRE? 2097152,1048576

:SYST:FRE? 3145728,1048576

Read the next 4 MB, the same data is read again:

:SYST:FRE? 4194304,1048576

:SYST:FRE? 5242880,1048576

:SYST:FRE? 6291456,1048576

:SYST:FRE? 7340032,1048576

:SYST:FWRite Address decimal, length decimal, data -> Write Flash memory

Only text data is written properly, eg. :SYST:FWR 4194048,8,ABCD1234

A byte with the value zero can be written only with tricks.

SCPI commands

TEST:ONE15

TEST:ONE16

TEST:ONE17

TEST:ONE18

TEST:ONE19

TEST:ONE20

TEST:ONE21

TEST:ONE22

TEST:ONE23

TEST:ONE24

TEST:ONE25

TEST:ONE26

TEST:ONE27

TEST:ONE28

TEST:ONE29

TEST:ONE30

TEST:ONE31

TEST:ONE32

TEST:ONE33

TEST:ONE34

TEST:ONE35

TEST:ONE36

TEST:ONE37

TEST:ONE38

TEST:ONE39

TEST:ONE40

TEST:ONE41

TEST:ONE42

TEST:ONE43

*CLS

*ESE *ESE?

*ESR?

*IDN?

*OPC *OPC?

*RST

*SRE *SRE?

*STB?

*TRG

*TST?

*WAI

ABORt ABOR

CALCulate CALC

BANDwidth BAND

NDB NDB?

RESult? RES?

MARKer MARK

AOFF

CPEak CPE

STATe STAT STATe? STAT?

DELTa DELT

SET

CENTer CENT

SPAN

FCOunt FCO

RESolution RES RESolution?

AUTO AUTO?

X?

FUNCtion FUNC FUNCtion? FUNC?

MAXimum MAX

LEFT

NEXT

RIGHt RIGH

MINimum MIN

MODE MODE?

PEAK

EXCursion EXC EXCursion? EXC?

SEARch SEAR

CF

THReshold THR THReshold? THR?

PTPeak PTP

RLEVel RLEV

STARt STAR

STEP

STOP

TABLe TABL

TRACe TRAC TRACe? TRAC?

TRCKing TRCK TRACking

X

CENTer? CENT?

POSition POS POSition? POS?

SPAN?

STARt? STAR?

STOP?

READout READ READout? READ?

Y?

VSRefl? VSR?

VSValue? VSV?

LLINe LLIN

ALL

DELete DEL

CONTrol CONT

DOMain DOM DOMain? DOM?

INTerpolate INT

TYPE TYPE?

DATA DATA?

MERGe MERG

FAIL FAIL?

RATIo? RATI?

RELFreq RELF RELFreq? RELF?

RELAmpt RELA RELAmpt? RELA?

NTData NTD

CALibration CAL

ACCUrate ACCU

SIGNal SIGN SIGNal? SIGN?

CONFigure CONF

ACPower ACP

CHPower CHP

CNRatio CNR

EBWidth EBW

HDISt HDIS

OBWidth OBW

SANalyzer SAN

TOI

TPOWer TPOW

PF

CONFigure? CONF?

COUPle COUP COUPle? COUP?

DISPlay DISP

AFUnction AFU

ANNotation ANN

CLOCk CLOC

BRIGtness BRIG BRIGtness? BRIG?

BRIGHtness BRIGH BRIGHtness? BRIGH?

FSCReen FSCR

ZSCReen ZSCR

SSCRenn SSCR

ENABle ENAB ENABle? ENAB?

MSGswitch MSG

WINdow WIN

GRATicule GRAT

GRID GRID?

Y

DLINe  DLIN DLINe? DLIN?

SCALe SCAL

PDIVision PDIV PDIVision? PDIV?

RLEVel? RLEV?

OFFSet OFFS OFFSet? OFFS?

SPACing SPAC SPACing? SPAC?

NRLevel NRL NRLevel? NRL?

NRPosition NRP NRPosition? NRP?

UKEY

FETCh FETC

ACPower? ACP?

LOWer? LOW?

UPPer? UPP?

MAIN?

CHPower? CHP?

DENSity? DENS?

CNRatio? CNR?

CARRier? CARR?

NOISe? NOIS?

EBWidth? EBW?

HARMonics HARM

AMPLitude AMPL AMPLitude? AMPL?

ALL?

DISTortion? DIST?

FREQuency FREQ FREQuency? FREQ?

FUNDamental? FUND?

OBWidth? OBW?

FERRor? FERR?

TOIntercept TOIntercept? TOI?

IP3?

TPOWer? TPOW?

FORMat FORM

BORDer BORD BORDer? BORD?

HCOPy HCOP

IMAGe IMAG

COLor COL

INVert INV INVert? INV?

PTIMe PTIM PTIMe? PTIM?

QUALity QUAL QUALity? QUAL?

FTYPe FTYP FTYPe? FTYP?

IMMediate IMM

PAGE

ORIentation ORI ORIentation? ORI?

PRINts PRIN PRINts? PRIN?

SIZE SIZE?

RESume

INITiate INIT

CONTinuous CONTinuous? CONT?

PAUSe PAUS

RESTart REST

INPut INP

IMPedance IMP IMPedance? IMP?

MMEMory MMEM

MOVE

LOAD

CORRection CORR

SETUp SETU

LIMit LIM

MTABle MTAB

STORe STOR

SCReen SCR

RESults

PTABle PTAB

DISK

INFomation? INF? INFOrmation? INFO?

OUTPut OUTP

SENSe  SENS

VIDeo VID VIDeo? VID?

RATio RAT RATio? RAT?

EMIFilter EMIF

CSET

DEMod DEM DEMod? DEM?

TIME TIME?

DETector DET

INCRement INCR INCRement? INCR?

UP

DOWN

FULL

PREVious PREV

ZIN

ZOUT

ZONE

RZONe RZON

LZONe LZON

SYSNThesis SYSNT SYSNThesis? SYSNT?

POWer POW

RF

ATTenuation ATT ATTenuation? ATT?

GAIN

MIXer MIX

RANGe RANG

UPPer UPP

ASCale ASC

ARANge ARAN

ATUNe  ATUN

SWEep SWE

COUNt COUN COUNt? COUN?

CURRent? CURR?

RULes RUL RULes? RUL?

AVERage AVER

TCONtrol TCON TCONtrol? TCON?

INTegration INTegration? INT?

ACHannel ACH ACHannel? ACH?

CSPacing CSP CSPacing? CSP?

NOISe NOIS

MAXHold MAXH

XDB XDB?

NUMBers NUMB NUMBers? NUMB?

PERCent PERC PERCent? PERC?

LLIMit LLIM LLIMit? LLIM?

RLIMit RLIM RLIMit? RLIM?

VSWR

RESet

FREFlect FREF

NREFlect NREF

SOURce SOUR

LEVel LEV

SWEep? SWE?

STORref

REF

ALC

STATus

OPERation OPER

CONDition? COND?

EVENt? EVEN?

PRESet PRES

QUEStionable QUES

SYSTem SYST

COMMunicate COMM

APORt APOR APORt? APOR?

GPIB

SELF

ADDRess ADDR ADDRess? ADDR?

LAN

DHCP

AUToip AUT

MANuip MAN

IP

ADDress ADD ADDress? ADD?

SUBMask SUBM SUBMask? SUBM?

GATeway GAT GATeway? GAT?

DNSServer DNSS DNSServer? DNSS?

DNAMe DNAM DNAMe? DNAM?

HNAMe HNAM HNAMe? HNAM?

PASSword PASS PASSword? PASS?

SERial SER

RECeive REC

BAUD BAUD?

PARity PAR PARity? PAR?

USB

CLASs CLAS CLASs? CLAS?

MESSage? MESS?

DATE DATE?

ERRor ERR

NEXT?

LANGuage LANG LANGuage? LANG?

LKEY LKEY?

OPTions? OPT?

PON

SAVE

SPEaker SPE

VOLume VOL VOLume? VOL?

BEEPer BEEP

LINemod LIN

TYPe TYP

USERkey USER

KEYCmd KEYC KEYCmd? KEYC?

CONFirm

VERSion? VERS?

FSWItch FSWI

TX

SWSTa? SWST?

SWSet SWS

FREad? FRE?

FWRite FWR

MATH

SUBTract SUBT

POINts? POIN?

SORT SORT?

A A?

B B?

CONSt CONS CONSt? CONS?

CLEar CLE

TRIGger TRIG

SEQuence SEQ

EXTernal EXT

SLOPe SLOP SLOPe? SLOP?

READy?

SOURce? SOUR?

LEVel? LEV?

TRIger TRI

UNIT

POWer? POW?

KEYBoard KEYB

KEYValue KEYV

SELFtest

KEY

SYS

SERVice SERV

LOG

DEVIce DEVI DEVIce? DEVI?

CPState CPS CPState? CPS?

PRIVate PRIV

AMPLref AMPLref?

FREQref FREQref?

LOdac LO LOdac? LO?

CLKDac CLKD CLKDac? CLKD?

ONDac OND ONDac? OND?

OFFDac OFFD OFFDac? OFFD?

WRITeflash WRIT

READflash?

WNVram WNV

RNVram? RNV?

TGCal TGC

TEMPeratue TEMP

GET?

COMPensate COMP

SNAPscr? SNAP?

TEST

FloodData FD FloodData? FD?

FACTory FACT

SN SN?

MODEL MODEL?

MAC MAC?

RTC RTC?

CALRtc CALR CALRtc? CALR?

AKEY AKEY?

SKEY SKEY?

SOFT

VER?

CTRL

DATAtype?

SYNC?

ESR?

CMD1

CMD2

CMD3

CMD4

CMD5

CMD6

CMD7

CMD8

CMD9

CMD10

CMD11

CMD12

CMD13

CMD14

CMD15

CMD16

CMD17

CMD18

CMD19

CMD20

CMD21

CMD22

CMD23

CMD24

CMD25

CMD26

CMD27

CMD28

CMD29

CMD30

CMD31

CMD32

CMD33

CMD34

CMD35

CMD36

CMD37

CMD38

CMD39

SCPI parameter

Rigol Technologies,%s,%s,%s

ERR

NULL

N/A

ALL

NONE

.cbl

.CBL

.CSV

CSV

TRC

TRACE1

TRACE2

TRACE3

MATH

MKR

OFF

ERR

.sta

BMP

JPEG

1LOCUR

1LOMUXOUT

1LOFREQ

1LOR0

1LOR1

1LOR2

1LOR3

1LOR4

1LORALL

1LOCH

1LOBW

2LOCUR

2LOMUXOUT

2LOFREQ

2LOR0

2LOR1

2LOR2

2LOR3

2LORALL

TGLOCUR

TGLOMUXOUT

TGLOFREQ

TGLOR0

TGLOR1

TGLOR2

TGLOR3

TGLOR4

TGLORALL

TGATT

DACA

DACB

LORATE

SPIPHASEEXE

SPIPHASEUP

SPIPHASESET

OFF

AM

FM

PM

POS

NEG

SAMP

NORA

RMS

VAV

ERR

NORM

ACC

EXP

REP

PEAK

AVER

QPEAK

SWE

FIX

IMM

EXT

VID

DBM

DBUV

DBMV

V

W

CLAS

MODE

DEF

DRAF

FINE

EXIF

PORT

LAND

A4

A5

A6

B5

FREQ

AMPL

DLM

DLL

WRIT

MAXH

MINH

POW

VIEW

BLANK

LOWERL

UPPERL

T1

T2

T3

T_MATH

A-B

A+CONST

A-CONST

NOIS

NDB

MAX

PAR

PER

TIME

ITIM

DELT

BAND

SPAN

TPOW

ACP

CHP

OBW

EBW

CNR

HD

TOI

PF

NULL

LOG

LIN

UNMEAS

PASS

FAIL

INT

LAN

USB

GPIB

115K

57.6K

38.4K

19.2K

9600

4800

2400

NONE

ODD

EVEN

ENGL

CHIN

JAPANESE

GERMAN

POLISH

PORTUGESE

KOREA

FACT

USER1

USER2

USER3

USER4

USER5

USER6

AUTO

TMC

PRIN

LAST

PRES

TOP

CENT

BOTT

SWAP

ASCII

REAL,32

PC

UDISK

RMT

USER

FACTORY

REPAIR

COLOR_EGG

ON

ANT

ANTENNA

CAB

CABLE

OTH

OTHER

EXPONENTIAL

REPEAT

NEGATIVE

NORMAL

POSITIVE

SAMPLE

VAVERAGE

QPE

ACCURACY

AVERAGE

FIXED

SWEEP

IMMEDIATE

EXTERNAL

VIDEO

CLASSIC

MODERN

DEFAULT

DRAFT

EXIFJPEG

LANDSCAPE

PORTRAIT

FREQUENCY

AMPLITUDE

DLMORE

DLLESS

WRITE

MAXHOLD

MINHOLD

VIDEOAVG

POWERAVG

BLAN

2

3

MATH

4

NOISE

POSITION

DELTA

MAXIMUM

PARAMETER

PERIOD

ITIME

LOGARITHMIC

LINEAR

INTERNAL

BOTTOM

CENTER

PRINTER

ENGLISH

CHINESE

JAPAN

JAP

GERM

POL

KOR

PRESET

ASC

REAL

SWAPPED

TRACE1

TRACE2

TRACE3

TRACE4

50

75

GHZ

MHZ

KHZ

HZ

PS

NS

US

MS

KS

S

SW1

SW2

SW3

SW4

SW5

Non-visible menu items. It is only the text available.

System -> Licence -> Install
     Del Option

Storage
     Cover file
     Reenter

Menu text english:

Frequency.Center Freq.Start Freq.Stop Freq.CF Step.Auto.Manual.Signal Track.On.

Off.Peak->CF.CF->Step.Span.Span.Full Span.Zero Span.Zoom In.ZoomOut.Last Span.

Amplitude.Auto Scale.Ref Level.Input Att..Auto.Manual.Scale/Div.Scale Type.Log.

Lin.Units.dBm.dBmV.dBuV.Volts.Watts.Ref Offset.Auto Range.RF Preamp.On.Off.

Corrections.Select.Off.Antenna.Cable.Other.User.Correction.On.Off.Edit.Point.

Frequency.Amplitude.Del Point.Freq Interp.Log.Lin.Delete.OK.Cancel.Corr Table.

On.Off.Corr View.All .Sel.MaxMixL.Input.50¦¸.75¦¸.BW/Det.RBW.Auto.Manual.VBW.

Auto.Manual.V/R Ratio.Det Type.Pos Peak.Neg Peak.Sample.Normal.RMS Avg.

Voltage Avg.Quasi-Peak.Filter Type.Gauss.EMI.Sweep/Trig.Time.Auto.Manual.

Auto SWT.Normal.Accy.Mode.Single.Cont.Single.Numbers.Trig Type.FreeRun.Video.

External.Trig Setup.Video.Trig Level.External.Edge.Pos.Neg.Auto .Demod.Demod.

Off.AM.FM.Demod Setup.AM.Earphone.On.Off.Volume.Demod Time.FM.Earphone.On.Off.

Volume.Demod Time.Trace/PF.Select Trace.1.2.3.Trace Type.Clear Write.Max Hold.

Min Hold.Video Avg.Power Avg.Freeze.Blank.Avg Times.Trace Math.Function.A-B.

A+Const.A-Const.A.T1.T2.T3.B.T1.T2.T3.Const.Operate.On.Off.Clear All.Pass/Fail.

Switch.On.Off.Setup.Limit.Upper.Lower.Test.On.Off.Edit.Limit.Upper.Lower.Point.

X Axis.Amplitude.Connected.Yes.No.Del Point.OK.Cancel.X Axis.Freq.Time.

Freq Interp.Log.Lin.Rel Setting.Rel Freq.On.Off.Rel Ampt.On.Off.Fail Stop.On.

Off.Del Limit.OK.Cancel.Beeper.On.Off.Restart.Pause.Resume.Meas Mode.Single.

Cont.Single.AVG Reset.TG.TG.On.Off.TG Level.TG Lvl Offset.Power Sweep.On.Off.

Power Range.Normalize.Stor Ref.Normalize.On.Off.Norm Ref Lvl.Norm Ref Pos.

Ref Trace.View.Blank.ALC.Auto.On.Off.Measure.VSWR.On.Off.Meas Fctn.Off.T-Power.

ACP.Chan Pwr.OBW.EBW.C/N Ratio.Harmo Dist.TOI.Restart.Pause.Resume.Meas Mode.

Single.Cont.Single.Meas Setup.T-Power.Avg Num.On.Off.Avg Mode.Exp.Repeat.

TP Type.Peak.Average.RMS.Start Line.Stop Line.ACP.Avg Num.On.Off.Avg Mode.Exp.

Repeat.Main CH BW.Adj CH BW.CH Spacing.Chan Power.Avg Num.On.Off.Avg Mode.Exp.

Repeat.Integ BW.CH PwrSpan.OBW.Avg Num.On.Off.Avg Mode.Exp.Repeat.Max Hold.On.

Off.OBW Span.Power Ratio.EBW.Avg Num.On.Off.Avg Mode.Exp.Repeat.Max Hold.On.Off.

EBW Span.EBW X dB.C/N Ratio.Avg Num.On.Off.Avg Mode.Exp.Repeat.Offset Freq.

Noise BW.Carrier BW.Harmo Dist.Avg Num.On.Off.Avg Mode.Exp.Repeat.NO.of Harmo.

Harmonic ST.Auto.Manual.TOI.Avg Num.On.Off.Avg Mode.Exp.Repeat.TOI Span.VSWR.

Reset.Cal Open.VSWR.Marker.1.2.3.4.Marker State.On.Off.Ref Lvl.Marker.

Select Mkr.1.2.3.4.Normal.Delta.Delta Pair.Ref.Delta.Span Pair.Span.Center.Off.

Mkr Trace.1.2.3.Math.Auto.Readout.Frequency.Period.¦¤Time.1/¦¤time.Mkr Table.On.

Off.All Off.Marker Fctn.Select Mkr.1.2.3.4.Noise Mkr.N dB BW.Function Off.

Freq Count.State.On.Off.Resolution.Auto.Manual.Marker->.Mkr->CF.Mkr->Step.

Mkr->Start.Mkr->Stop.Mkr->Ref.Mkr¦¤->CF.Mkr¦¤->Span.Peak.Next Peak.Peak Right.

Peak Left.Min Search.Peak Peak.Cont Peak.On.Off.Search Para.Pk Excursn.

Pk Thresh.Peak Search.Max.param.Peak Table.State.On.Off.Peak Sort.Freq.Ampl.

Pk Readout.Normal.>DL.<DL.Preset.System.Language.English.ÖÐÎÄ.ÈÕ±¾ÕZ.Portugu¨ºs.

Deutsch.Polski.???.Reset.Power On.Last.Preset.Preset Type.Factory.User1.User2.

User3.User4.User5.User6.Save Preset.OK.Cancel.Calibrate.Cal Now.Self-Cal.On.Off.

Acc Cal.I/O Setting.Remote I/O.Off.LAN.USB.GPIB.LAN.Reset.OK.Cancel.Config.OK.

Cancel.DHCP.On.Off.Auto-IP.On.Off.Manual-IP.On.Off.IP.IP Address.Subnet Mask.

Gateway.DNS.USB.Dev Class.AutoConfig.TMC.Printer.Dev Addr.GPIB.Display.

DisplayLine.On.Off.Active Fctn.Top.Center.Bottom.Graticule.Scr State.On.Off.

Brightness.UserKey.On.Off.Msg Switch.On.Off.Work Setting.Front Switch.On.Off.

Line Mode.Factory.User1.User2.User3.User4.User5.User6.UserKey Set.On.Off.

Coupl Param.Information.System Info.System Msg.Self-Test.Screen Test.Key Test.

Time/Date.Time/Date.On.Off.Set Time.Set Date.License.Option Info.License Info.

Install.OK.Cancel.Del Option.TX1000.Sw1.On.Off.Sw2.On.Off.Sw3.On.Off.Sw4.On.Off.

Sw5.On.Off.Service.Calibration.Cal Switch.On.Off.Int Ref Signal.On.Off.

Clr Int Cal.Save Cal.Ref DAC.Cal DAC.LO DAC.TG Cal.On.Off.TG Att.Clk DAC.

Off DAC.On DAC.Log Manage.Log.On.Off.OutPut Dev.PC.U-Disk.Rmt-Monitor.CP state.

Signal Ctrl.Ph Noise Opt.On.Off.Storage.File Type.All.Setup.State.Trace.

Corrections.Limit.Measure.Mkr Table.Peak Table.Format.BIN.CSV.File Source.T1.

T2.T3.Trace Math.Trace All.Browser.Dir.File.Save.OK.Cover file.Reenter.Save.

Cancel.Cancel.Expand Dir.Collapse Dir.Recall.Rename.OK.Cancel.Delete.OK.Cancel.

Copy.Copy From.Copy To.Replace File.Cancel.Apply To.User1.User2.User3.User4.

User5.User6.Browser.Dir.File.Expand Dir.Collapse Dir.Create Dir.OK.Cancel.

Disk Info.Name Prefix.Prefix Switch.On.Off.Edit Prefix.OK.Cancel.Sys Update.

Print .Print Setup.Print.Resume.Cancel.Orientation.Portr.Landsc.Page Size.

Default.A4.A5.A6.B5.Inverted.On.Off.Palette.Gray.Color.Copies.Date Prints.On.

Off.Qualities.Default.Normal.Draft.Fine.FileType.Default.Exif/JPEG.BMP.Help.

User Key.Marker.Marker.Delta Marker.Marker Ref.Marker Delta.Marker Span.

Marker Center.Unit.U_Freq.GHz.MHz.kHz.Hz.U_Amp.U_dBm.dBm.-dBm.mV.uV.U_dBmV.dBmV.

-dBmV.mV.uV.U_dBuV.dBuV.-dBuV.mV.uV.U_Volts.V.mV.uV.nV.dBm.U_Watts.W.mW.uW.nW.

dBm.U_dB.dB.U_Time.ks.s.ms.us.ns.ps.U_Enter.Enter.U_Percent.%.LIST.L_T_POWER.

T-POWER.TP Type.L_ACP.Main CH Pwr.Upper.Lower.Adjacent Channel Power.L_CHP.

Channel Power.Power Spectral Density.L_OBW.Occupied BandWidth.

Transmit Freq Error.L_EBW.Emission BandWidth.L_CNR.Carrier Power.Noise Power.

CNRatior.L_HD.THD.Freq.Amp.Real Harmonics.Sweep Time.Harmonic Distortion.L_TOI.

dBc.Intercept.Base Lower.Base Upper.3rd Order Lower.3rd Order Upper.Pass.

P/F Ratio.L_MARKER.Marker.Trace.Type.X Axis.MarkerTable.Refl Coefficient.

Return Loss.L_PEAKER.Peak.L_CORR.Num.L_LAN.IP Setting.L_INFO_MSG.Msg ID.

Msg Content.Time.Line Mode: .Esc to exit..Invalid sweep Data.L_DISK_INFO.

Disk name.Type.Local disk.Mobile disk.File system.Used space.Free space.

Total space.Byte.Opt Install.Offical.Trial.Invalid.None.Forever

\SDRAM_FWv00.01.09.00.07_with-TG-TGCal-Options\dsa815_sdram.bin

........

........

........  SCPI parameter, commands

00470db0  52 69 67 6f 6c 20 54 65 63 68 6e 6f 6c 6f 67 69  Rigol Technologi

00470dc0  65 73 2c 25 73 2c 25 73 2c 25 73 00 24 21 0c 00  es,%s,%s,%s.$!..

........  .. .. .. .. .. .. .. .. .. .. .. .. .. .. .. ..  ................

00470f50  31 4c 4f 43 55 52 00 00 31 4c 4f 4d 55 58 4f 55  1LOCUR..1LOMUXOU

00470f60  54 00 00 00 31 4c 4f 46 52 45 51 00 31 4c 4f 52  T...1LOFREQ.1LOR

........  .. .. .. .. .. .. .. .. .. .. .. .. .. .. .. ..  ................

00471110  4f 46 46 00 41 4d 00 00 46 4d 00 00 50 4d 00 00  OFF.AM..FM..PM..

00471120  50 4f 53 00 4e 45 47 00 53 41 4d 50 00 00 00 00  POS.NEG.SAMP....

........  .. .. .. .. .. .. .. .. .. .. .. .. .. .. .. ..  ................

00473240  35 00 00 00 43 4d 44 33 36 00 00 00 43 4d 44 33  5...CMD36...CMD3

00473250  37 00 00 00 43 4d 44 33 38 00 00 00 43 4d 44 33  7...CMD38...CMD3

00473260  39 00 00 00 9c 19 47 00 9c 19 47 00 ff ff 01 00  9...œ.G.œ.G.ÿÿ..

........

........

........  Menu text english

00603000  46 72 65 71 75 65 6e 63 79 00 43 65 6e 74 65 72  Frequency.Center

00603010  20 46 72 65 71 00 53 74 61 72 74 20 46 72 65 71   Freq.Start Freq

00603020  00 53 74 6f 70 20 46 72 65 71 00 43 46 20 53 74  .Stop Freq.CF St

00603030  65 70 00 41 75 74 6f 00 4d 61 6e 75 61 6c 00 53  ep.Auto.Manual.S

00603040  69 67 6e 61 6c 20 54 72 61 63 6b 00 4f 6e 00 4f  ignal Track.On.O

00603050  66 66 00 50 65 61 6b 2d 3e 43 46 00 43 46 2d 3e  ff.Peak->CF.CF->

00603060  53 74 65 70 00 53 70 61 6e 00 53 70 61 6e 00 46  Step.Span.Span.F

........  .. .. .. .. .. .. .. .. .. .. .. .. .. .. .. ..  ................

........  .. .. .. .. .. .. .. .. .. .. .. .. .. .. .. ..  ................

........  .. .. .. .. .. .. .. .. .. .. .. .. .. .. .. ..  ................

........

........

........  Key ScanCode

01883780  1e 30 00 00 12 00 00 00 46 52 45 51 00 00 00 00  .0......FREQ....

01883790  00 00 00 00 22 00 00 00 53 50 41 4e 00 00 00 00  ...."...SPAN....

018837a0  00 00 00 00 32 00 00 00 41 4d 50 00 00 00 00 00  ....2...AMP.....

018837b0  00 00 00 00 52 00 00 00 42 57 00 00 00 00 00 00  ....R...BW......

018837c0  00 00 00 00 62 00 00 00 53 57 45 45 50 00 00 00  ....b...SWEEP...

018837d0  00 00 00 00 a2 00 00 00 54 55 4e 45 00 00 00 00  ....¢...TUNE....

018837e0  00 00 00 00 c2 00 00 00 44 45 4d 4f 44 00 00 00  ....Â...DEMOD...

018837f0  00 00 00 00 d2 00 00 00 54 52 41 43 45 00 00 00  ....Ò...TRACE...

01883800  00 00 00 00 e2 00 00 00 54 47 00 00 00 00 00 00  ....â...TG......

01883810  00 00 00 00 f2 00 00 00 4d 45 41 53 00 00 00 00  ....ò...MEAS....

01883820  00 00 00 00 02 01 00 00 4d 45 41 53 53 45 54 00  ........MEASSET.

01883830  00 00 00 00 12 01 00 00 4d 41 52 4b 00 00 00 00  ........MARK....

01883840  00 00 00 00 22 01 00 00 4d 41 52 4b 46 55 4e 43  ...."...MARKFUNC

01883850  00 00 00 00 32 01 00 00 4d 41 52 4b 54 4f 00 00  ....2...MARKTO..

01883860  00 00 00 00 42 01 00 00 50 45 41 4b 00 00 00 00  ....B...PEAK....

01883870  00 00 00 00 72 01 00 00 50 52 45 53 45 54 00 00  ....r...PRESET..

01883880  00 00 00 00 82 01 00 00 53 59 53 54 45 4d 00 00  ....‚...SYSTEM..

01883890  00 00 00 00 92 01 00 00 53 54 4f 52 41 47 45 00  ....’...STORAGE.

018838a0  00 00 00 00 b2 01 00 00 50 52 49 4e 54 00 00 00  ....²...PRINT...

018838b0  00 00 00 00 c2 01 00 00 50 52 49 4e 54 53 45 54  ....Â...PRINTSET

018838c0  55 50 00 00 d2 01 00 00 48 45 4c 50 00 00 00 00  UP..Ò...HELP....

018838d0  00 00 00 00 02 03 00 00 46 31 00 00 00 00 00 00  ........F1......

018838e0  00 00 00 00 12 03 00 00 46 32 00 00 00 00 00 00  ........F2......

018838f0  00 00 00 00 22 03 00 00 46 33 00 00 00 00 00 00  ...."...F3......

01883900  00 00 00 00 32 03 00 00 46 34 00 00 00 00 00 00  ....2...F4......

01883910  00 00 00 00 42 03 00 00 46 35 00 00 00 00 00 00  ....B...F5......

01883920  00 00 00 00 52 03 00 00 46 36 00 00 00 00 00 00  ....R...F6......

01883930  00 00 00 00 62 03 00 00 46 37 00 00 00 00 00 00  ....b...F7......

01883940  00 00 00 00 72 03 00 00 52 45 54 55 52 4e 00 00  ....r...RETURN..

01883950  00 00 00 00 82 03 00 00 50 41 47 45 44 4f 57 4e  ....‚...PAGEDOWN

01883960  00 00 00 00 12 04 00 00 55 50 00 00 00 00 00 00  ........UP......

01883970  00 00 00 00 02 04 00 00 44 4f 57 4e 00 00 00 00  ........DOWN....

01883980  00 00 00 00 42 04 00 00 4b 4e 4f 42 4c 00 00 00  ....B...KNOBL...

01883990  00 00 00 00 52 04 00 00 4b 4e 4f 42 52 00 00 00  ....R...KNOBR...

018839a0  00 00 00 00 02 05 00 00 44 49 47 49 54 30 00 00  ........DIGIT0..

018839b0  00 00 00 00 12 05 00 00 44 49 47 49 54 31 00 00  ........DIGIT1..

018839c0  00 00 00 00 22 05 00 00 44 49 47 49 54 32 00 00  ...."...DIGIT2..

018839d0  00 00 00 00 32 05 00 00 44 49 47 49 54 33 00 00  ....2...DIGIT3..

018839e0  00 00 00 00 42 05 00 00 44 49 47 49 54 34 00 00  ....B...DIGIT4..

018839f0  00 00 00 00 52 05 00 00 44 49 47 49 54 35 00 00  ....R...DIGIT5..

01883a00  00 00 00 00 62 05 00 00 44 49 47 49 54 36 00 00  ....b...DIGIT6..

01883a10  00 00 00 00 72 05 00 00 44 49 47 49 54 37 00 00  ....r...DIGIT7..

01883a20  00 00 00 00 82 05 00 00 44 49 47 49 54 38 00 00  ....‚...DIGIT8..

01883a30  00 00 00 00 92 05 00 00 44 49 47 49 54 39 00 00  ....’...DIGIT9..

01883a40  00 00 00 00 a2 05 00 00 44 4f 54 00 00 00 00 00  ....¢...DOT.....

01883a50  00 00 00 00 c2 05 00 00 42 41 43 4b 00 00 00 00  ....Â...BACK....

01883a60  00 00 00 00 d2 05 00 00 45 53 43 00 00 00 00 00  ....Ò...ESC.....

01883a70  00 00 00 00 e2 05 00 00 45 4e 54 45 52 00 00 00  ....â...ENTER...

........

........

........  Factory Boot ModeMaintenance ModeUnknown Mode 1Unknown Mode 2User Mode

0188aea0  01 00 00 00 53 45 43 52 45 54 00 00 31 32 33 34  ....SECRET..1234

0188aeb0  35 30 30 30 31 32 33 34 35 36 00 00 00 00 00 00  5000123456......

0188aec0  00 00 00 00 00 00 00 00 00 00 00 00 00 00 c9 f0  ..............Éð

0188aed0  25 2f 00 00 30 30 30 31 00 30 30 30 32 00 30 30  %/..0001.0002.00

0188aee0  30 33 00 30 30 30 34 00 30 30 30 35 00 00 00 00  03.0004.0005....

0188aef0  06 01 69 ad 68 0c b2 01 14 00 00 00 00 00 00 00  ..ih.².........

0188af00  06 00 3c ad 7c 0c b2 01 54 00 00 00 00 00 00 00  ..<|.².T.......

0188af10  00 02 00 00 09 03 05 01 09 08 00 00 09 01 07 08  ................

0188af20  41 73 54 55 56 57 58 4e 4f 50 44 36 42 43 6a 4b  AsTBVWYNOPD6BCjK

0188af30  4c 4d 63 6b 6c 6d 6e 45 46 76 77 78 79 7a 74 6f  LMcklmnEFvsxyzto

0188af40  70 34 35 47 71 75 52 53 59 5a 61 62 72 30 31 32  p45Gq8RSYZabr012

0188af50  33 48 49 4a 64 65 66 67 68 69 51 37 38 39 23 2e  3HIJdeaghiQ789#.

0188af60  00 00 00 00 06 01 a5 ad d0 0c b2 01 60 00 00 00  ......¥Ð.².`...

0188af70  00 00 00 00 06 00 54 ad 30 0d b2 01 70 01 00 00  ......T0.².p...

0188af80  00 00 00 00 05 00 00 00 01 00 00 00 02 00 00 00  ................

0188af90  03 00 00 00 04 00 00 00 05 00 00 00 52 00 00 00  ............R...

0188afa0  62 00 00 00 12 01 00 00 e2 00 00 00 f2 00 00 00  b.......â...ò...

0188afb0  02 01 00 00 c2 00 00 00 d2 00 00 00 e2 00 00 00  ....Â...Ò...â...

0188afc0  22 01 00 00 02 01 00 00 82 01 00 00 c2 01 00 00  ".......‚...Â...

0188afd0  92 01 00 00 52 00 00 00 12 01 00 00 d2 00 00 00  ’...R.......Ò...

0188afe0  e2 00 00 00 22 01 00 00 62 00 00 00 e2 00 00 00  â..."...b...â...

0188aff0  02 01 00 00 c2 01 00 00 e2 00 00 00 f2 00 00 00  ....Â...â...ò...

0188b000  82 01 00 00 e2 00 00 00 c2 00 00 00 92 01 00 00  ‚...â...Â...’...

0188b010  e2 05 00 00 d2 05 00 00 c2 05 00 00 12 00 00 00  â...Ò...Â.......

0188b020  12 00 00 00 12 00 00 00 46 41 43 54 4f 52 59 4d  ........FACTORYM

0188b030  4f 44 45 45 4e 54 52 41 4e 43 45 00 52 45 50 41  ODEENTRANCE.REPA

0188b040  49 52 4d 4f 44 45 45 4e 54 52 41 4e 43 45 00 00  IRMODEENTRANCE..

0188b050  45 58 49 54 43 55 52 52 45 4e 54 4d 4f 44 45 00  EXITCURRENTMODE.

0188b060  33 35 4f 38 4f 32 32 38 4f 4c 4f 38 4c 4e 4d 4e  35O8O238OLO8LNMN

0188b070  39 36 39 30 32 31 37 39 36 33 4d 4e 32 4f 39 4f  9690217963MN2O9O

0188b080  4e 34 36 4f 39 50 30 35 39 37 35 4b 30 30 34 37  N46O9P05975K0047

0188b090  33 4c 35 37 38 31 34 31 4b 39 36 32 39 35 36 31  3L568141K9629561

0188b0a0  31 30 32 32 4c 38 4e 33 37 34 36 33 39 35 32 34  1022K8N374639524

0188b0b0  30 4c 32 33 32 33 4f 31 36 30 32 33 32 34 34 4b  0L2323E16023244K

0188b0c0  34 31 32 34 31 34 38 37 30 30 33 4c 35 35 50 37  41241487013L55P7

0188b0d0  4b 36 4d 4c 4b 4b 36 33 30 36 34 39 4d 39 33 31  K6MLKK633649M931

0188b0e0  4b 4f 34 4c 30 35 36 37 32 38 4e 33 39 39 50 4f  KO4L056729N399PO

0188b0f0  00 00 00 00 06 01 9d ad a0 0e b2 01 10 3b 01 00  ...... .²..;..

........

........  It's all twice.

........

........  SCPI parameter, commands

0186a490  80 00 00 00 00 00 00 19 52 69 67 6f 6c 20 54 65  €.......Rigol Te

0186a4a0  63 68 6e 6f 6c 6f 67 69 65 73 2c 25 73 2c 25 73  chnologies,%s,%s

0186a4b0  2c 25 73 00 24 21 0c 00 26 21 0c 00 48 21 0c 00  ,%s.$!..&!..H!..

........

........

........

........  Key ScanCode

01b18e50  00 00 00 00 1e 30 00 00 12 00 00 00 46 52 45 51  .....0......FREQ

01b18e60  00 00 00 00 00 00 00 00 22 00 00 00 53 50 41 4e  ........"...SPAN

01b18e70  00 00 00 00 00 00 00 00 32 00 00 00 41 4d 50 00  ........2...AMP.

........

........

........  Factory Boot ModeMaintenance ModeUnknown Mode 1Unknown Mode 2User Mode

01b20c10  19 40 34 40 ff 11 00 00 01 00 00 00 53 45 43 52  .@4@ÿ.......SECR

01b20c20  45 54 00 00 31 32 33 34 35 30 30 30 31 32 33 34  ET..123450001234

01b20c30  35 36 00 00 a0 42 65 21 e6 97 01 00 d2 19 92 a9  56.. Be!æ—..Ò.’©

........

........

........

Hidden Modes - Graphic

Correction file

\testcal4.cbl - 322 Bytes

00000000  39 93 3c 01 00 00 ff 11 00 00 43 42 4c 46 38 01  9“<...ÿ...CBLF8.

00000010  00 00 14 00 64 00 c0 00 00 00 00 00 00 00 01 00  ....d.À.........

00000020  00 00 43 4f 52 52 c0 00 00 00 4e 4f 4e 45 00 00  ..CORRÀ...NONE..

00000030  00 00 4e 4f 4e 45 00 00 00 00 4e 4f 4e 45 00 00  ..NONE....NONE..

00000040  00 00 4e 4f 4e 45 00 00 00 00 4e 4f 4e 45 00 00  ..NONE....NONE..

00000050  00 00 4e 4f 4e 45 00 00 00 00 4e 4f 4e 45 00 00  ..NONE....NONE..

00000060  00 00 4e 4f 4e 45 00 00 00 00 4e 4f 4e 45 00 00  ..NONE....NONE..

00000070  00 00 30 30 2e 30 31 2e 30 39 2e 30 30 2e 30 37  ..00.01.09.00.07

00000080  00 00 00 00 a0 c0 80 77 8e 06 00 00 00 00 00 00  .... À€wŽ.......

00000090  80 3f 00 c2 eb 0b 00 00 00 00 00 00 80 3f 00 a3  €?.Âë.......€?.£

000000a0  e1 11 00 00 00 00 00 00 80 3f 00 84 d7 17 00 00  á.......€?.„×...

000000b0  00 00 00 00 80 3f 00 65 cd 1d 00 00 00 00 00 00  ....€?.eÍ.......

000000c0  80 3f 00 46 c3 23 00 00 00 00 00 00 00 40 00 27  €?.FÃ#.......@.'

000000d0  b9 29 00 00 00 00 00 00 e0 40 00 08 af 2f 00 00  ¹)......à@..¯/..

000000e0  00 00 00 00 00 40 00 e9 a4 35 00 00 00 00 00 00  .....@.é¤5......

000000f0  00 40 00 ca 9a 3b 00 00 00 00 00 00 00 40 00 ab  .@.Êš;.......@.«

00000100  90 41 00 00 00 00 00 00 40 40 00 8c 86 47 00 00  A......@@.Œ†G..

00000110  00 00 00 00 40 40 00 6d 7c 4d 00 00 00 00 00 00  ....@@.m|M......

00000120  40 40 00 4e 72 53 00 00 00 00 00 00 40 40 80 3e  @@.NrS......@@€>

00000130  6d 56 00 00 00 00 00 00 80 c0 c0 ec 58 59 00 00  mV......€ÀÀìXY..

00000140  00 00                                            ..

testcal4.cbl - CorrTable

Point 1    110 MHz   -5 dB

Point 2    200 MHz    1 dB

Point 3    300 MHz    1 dB

Point 4    400 MHz    1 dB

Point 5    500 MHz    1 dB

Point 6    600 MHz    1 dB

Point 7    700 MHz    2 dB

Point 8    800 MHz    7 dB

Point 9    900 MHz    2 dB

Point 10  1000 MHz    2 dB

Point 11  1100 MHz    2 dB

Point 12  1200 MHz    3 dB

Point 13  1300 MHz    3 dB

Point 14  1400 MHz    3 dB

Point 15  1450 MHz    3 dB

Point 16  1499 MHz   -4 dB

Address 00000000  39 93 3c 01 = Is usually different per file, but not always.

Address 0000000A  43 42 4c 46 = Format identifier?

Address 0000000E  38 01 00 00 = 0x00000138 = 312 + 10 = 322 Bytes File size.

Address 00000016  c0 00 00 00 = 0x000000C0 = 192 Bytes Points data.

Address 00000022  43 4f 52 52 = „CORR“ = Correction data, Format identifier.

Address 00000026  c0 00 00 00 = 0x000000C0 = 192 Bytes Points data.

Address 00000072  30 30 2e 30 31 2e 30 39 2e 30 30 2e 30 37 = Firmware version.

Point 1    110 MHz   -5 dB  - 00000082  00 00 a0 c0 80 77 8e 06 00 00 00 00

Point 2    200 MHz    1 dB  - 0000008E  00 00 80 3f 00 c2 eb 0b 00 00 00 00

.......

.......

.......

Point 1 - Address 00000082

Float 32 bit  00 00 a0 c0 = 0xC0A00000 = -5.0E0 = -5 dB

Integer       80 77 8e 06 = 0x068E7780 = 110000000 = 110 MHz

Unknown       00 00 00 00

Point 2 - Address 0000008E

Float 32 bit  00 00 80 3f = 0x3F800000 = 1.0E0 = 1 dB

Integer       00 c2 eb 0b = 0x0BEBC200 = 200000000 = 200 MHz

Unknown       00 00 00 00

.......

.......

.......