LAB-View-Drivers for Devices with integrated INT2E Interface card

Available drivers:	EL-Series; ELD-Series; LA-Series, LD-Series, ZPS-Series,

ELA-Series

Components of the Lab-View-Drivers:

All Virtual Instruments (VI) needed for using Lab-View to control the devices.

System requirements:

Operating System: 	Windows 95/98/2000/NT

Software:		National-Instruments Lab-View

National-Instruments NI-VISA-Driver (Runtime engine)

(http://www.ni.com/support/visasupp.htm)

Components of the Driver for Series EL / ELD:

EL_Initialize.VI	Initialization of the interface and the electronic load.

EL_Reset.VI	Resets the INT2E.

EL_CLOSE.VI	Closes the NI-VISA interface.

EL_READMEAS.VI	Reads the actual values (voltage / current / power) and status information.

EL_WRITE2.VI	Writes the set values (voltage / current / power / conductance) to the electronic load.

EL_SRE.VI	Configuration of the INT2E enable registers.

(Provides mask data for enable registers: SRE, ESE, PPE)

EL_ID2MAXVAL.VI	Converts the IDN-String into numerical data and strings.

(Used for configuration of the Driver: Type, serial number, max. power, max. current, max voltage)

EL_READEREG.VI	Reads all registers that are not cleared after reading.

(STB, ESE, SRE, PRE)

EL_READREG.VI	Reads all registers. (Note: ESR is cleared after reading)

(STB, ESE, SRE, PRE, ESR)

EL_VISA_R.VI		Reads data from the interface.

EL_VISA_W.VI		Writes data into the interface.

EL_MAIN1.VI	Demo-VI (Test program to control all functions of the electronic load (Series EL / ELD)). It can be configured for IEEE488 (GPIB) or RS232 communication.

Controlling of Devices with integrated INT2E:

There are two ways to control the devices:

1.	No special driver used:

To control GPIB-Devices command strings can be send to control the devices. The user must create the command strings and send them to the device using the standard GPIB tools of LabView.

2. Using the LabView drivers:

The driver VIs (virtual instruments) are a set of software routines that control a programmable instrument. These easy to use routines simplify instrument control and reduce test  program development time. The VIs convert the LabView actions or  events into command strings and eliminate the need to learn the device specific command set and protocol.

To control our devices with Lab-View the driver VIs are used in the same way the standard VIs delivered with LabView are used. The driver VIs only can be executed under LabView. To use the driver LabView must be installed on each computer used to execute the created procedures (e.g. DEVICE_TEST.VI).

The Lab-View-Application-Builder is used to create executable Files that run independently (e.g. DEVICE_TEST.EXE).

Installation:

1. 	Copy the directory EL_DRIVER on disk 1 (including all subdirectories) onto your 	hard drive.

Copy the Demo-Program on disk 2 onto your hard drive into the directory: 	EL_DRIVER\VIs.

2. 	Ensure the VISA-Driver is installed.

If not installed: Download the NI-VISA 2.60 from the NI homepage.

(http://www.ni.com/support/visasupp.htm)

Install the NI-VISA driver

Running the test program:

1. 	Configure the interface in the electronic load using the dip-switch on the rear side of the electronic load.

1.1	RS232-Mode:	 9600Baud, no parity, 8 bit data, 1 stop bit

0

1

SW1	   1       2      3      4       5       6      7      8

Use a special 0-modem-cable

0-Modem-Cable (Manufacturer: Boeder)

| X1 |  | X2 | (SUB-D 9pin female) |
| --- | --- | --- | --- |
| Pin |  | Pin |  |
| 1 | <-> | 7 | DCD<->RTS |
| 2 | <-> | 3 | RxD<->TxD |
| 3 | <-> | 2 | TxD<->RxD |
| 4 | <-> | 6,8 | DTR<->DSR,CTS |
| 5 | <-> | 5 | GND<->GND |
| 6,8 | <-> | 4 | DSR,CTS<->DTR |
| 7 | <-> | 1 | RTS<->DCD |
| 9 | <-> | 9 | RI<->RI |

1.2	IEEE488-Mode: Default address: 4

0

1

SW1	   1       2      3      4       5       6      7      8

3.	Connect the computer and the electronic load using a 0-modem cable or

an IEEE488 cable.

4.	Switch on power.

5.	Load EL_MAIN1.VI (if necessary set new path for sub VIs and safe the VI)

6.	Choose VISA Resource Name:

IEEE488:	ASRL10

COM1: 		ASRL1

COM2: 		ASRL2

7.	Start EL_Main1.VI