UT2000/3000 Series Digital Storage Oscilloscopes RS232 Interface measurement parameters shows

1. Data Format

Divided into two parts: the first part of the data header, 7 bytes; second part of the measured parameters, the 20 measured parameters, measurement parameters for each accounting for 7 bytes, a total of 140 bytes; the two parts of a total of 147 characters section.

# A. Data header

| Serial number | 1 | 2 | 3 | 4 | 5 | 6 | 7 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Content | 0xAA | 0x55 | 0: CH1 1: CH2 | 0 | 0 | 0 | 0 |

# B. Measurement parameters

This part covers a total of 20 measured parameters, each parameter of each measure 7 bytes, now the format for each measurement parameter list is as follows:

| Serial number | 1 | 2 | 3 | 4 | 5 | 6 | 7 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Content | Measurements |  |  |  | Units of measurement |  |  |

Measured by the four bytes in the procedures of these four bytes when converted to a float.
Measurement unit consists of three bytes, three bytes of type character, each value corresponds to the corresponding ASCII code.

For example:

1). Receive a set of parameter values measured 7 bytes are:

0x00, 0x00, 0xA0, 0x40,  0x4D, 0x48, 0x7A

Measurements are converted to float type      Units, corresponding to ASCII code “MHz”

After the type is 5.00

2). Receive a set of parameter values measured 7 bytes are:

0x00, 0x00, 0x48, 0x43,  0x6E, 0x73, 0x00

Measurements are converted to float type         Units, corresponding to ASCII code “ns”

After the type is 200.00                     The last byte is meaningless

# C. Parameter array

PC sent a request to send parameters, send a one-time oscilloscope channel all the parameters under the corresponding total of 147 bytes, this 147 bytes in addition to the first 7 bytes of data header, the following 140 bytes are corresponding to the measured parameters, a total of 20 bytes each accounted for 7. The 20 parameters of the ranking are:

| Serial number | Share bytes | Content | Serial number | Share bytes | Content |
| --- | --- | --- | --- | --- | --- |
| 1 | 7 | Frequency | 11 | 7 | Mean |
| 2 | 7 | Cycle | 12 | 7 | Peak to peak |
| 3 | 7 | Rise time | 13 | 7 | RMS |
| 4 | 7 | Fall time | 14 | 7 | Top value |
| 5 | 7 | Are pulse width | 15 | 7 | Bottom value |
| 6 | 7 | Negative pulse | 16 | 7 | Median |
| 7 | 7 | Overshoot | 17 | 7 | Maximum |
| 8 | 7 | Pre-washed | 18 | 7 | Minimum |
| 9 | 7 | Positive duty cycle | 19 | 7 | Width |
| 10 | 7 | Negative duty cycle | 20 | 7 | Delay |

2. Control flow

Serial port initialization (open serial port, set the baud rate, stop bits, parity bit, data bits)

Successful?

Send measurement parameters command

Receiving parameter data

Return

Serial port initialization

Configure serial port parameters

- Baud Rate: 4800bps (only one baud rate)

- parity bit: none

- Data bits: 8

- Stop bit: 1

3. Control commands

2: 0xF9 and 0xFA.

CH1 transmission parameters, send control commands 0xF9;

CH2 transmission parameters, send control commands 0xFA;

4. Serial data receiving parameter reference C program

//////////////////////////////////////////////////////////////////////////////////////////////////////////////////

float dpara[20];   //// Used to store measurements

char paraunit[20][3]; //// Corresponding measured values for the storage unit

char  data1[147];

union

{

float *pfloat;

char *pchar;

}pdata;

ComWrtByte (1,0xF9);  //// CH1 through the serial port transfer command parameter

for(i=0;i<20000;i++) ;  /// Delay

ComRd (ComN, data1, 147); // Reading the serial port 147 bytes, and stored in an array of data1 in

if(data1[0]==0xaa  &&  data1[1]==0x55)/// Determine whether valid data

{

ch = data1[2];  ///// The upload of data corresponding to the channel parameters

pdata.pchar = data1 + 7;//// Parameter data area pointer

for (i=0; i<20; i++)////// Data classification caching

{

dpara[i] = *pdata.pfloat++;/// Store data

paraunit[i][0] = *pdata.pchar++;  ///// Storage Units

paraunit[i][1] = *pdata.pchar++;  ///// Storage Units

paraunit[i][2] = *pdata.pchar++;  ///// Storage Units

}

}

//////////////////////////////////////////////////////////////////////////////////////////////////////////////////